# REPORT TO THE AAPM RADIATION THERAPY COMMITTEE

Report No. 115		March 1, 2003 - July 31, 2003

## Personnel

There have been no changes in key personnel since the previous report.

## GRANT ACTIVITIES

We are presently in the fourth year of the current five year cycle.  Our application for competing renewal will be due at NCI on February 1, 2004.  A mock site visit will be scheduled for later this fall, and the NCI Site Visit is expected to be held in May or June,  2004.  We have began work on the application, and anticipate having a draft at the time of the ASTRO RTC Meeting.

The Advanced Technology for Radiation Therapy Grant to Washington University in St. Louis continues and has just completed its first year.  The RPC is a subcontractor to this grant.  The ATC met face to face in Montreal, Quebec in June at the time of the RTOG  Summer Meeting. Monthly telephone conference calls are held, and a short meeting of the ATC subcontractors is being held at this summer meeting in San Diego.

Recent activities of the RPC with regard to the ATC include integration of ITC-RCET software tools into our credentialing and chart review processes.  We have been using these tools as part of the IMRT and prostate brachytherapy credentialing programs.  We anticipate using the tools in the near future for data submissions to GOG and NCCTG clinical trials.  The RPC has contributed to discussions regarding a benchmark plan for IMRT credentialing developed by QARC.  A number of revisions to the benchmark plan were recently suggested by the RPC.  A benchmark plan for 3D treatment planning credentialing was recently posted on the RPC website.  This benchmark plan is based on one developed by QARC but has been adapted specifically for use by NCCTG institutions.

We recently discussed posting on the RPC website the status of each participating institution's credentials.  Further action on this proposal has been postponed, because of concerns about restricting access to the information.

The COMS subcontract continues, although the final technical meetings of the protocol will be held this fall.

The RPC also recently completed a contract with AMGEN Corporation to support a study of total body irradiation.  The RPC has provided the data collected to AMGEN.

BUDGET STATUS

As reported earlier, we received approval from the NCI to raise our annual fees to $500 per institution per year.  The increased fee went into effect when invoices were mailed to institutions last October.  We believe our budget is in good shape.

We were recently audited by the MDACC Internal Audit Department.  Preliminary indications are that everything was found to be in order

COMMUNICATIONS

# We continue to add items to the RPC webpage, and are now considering significant changes to improve its functionality.  We continue to post monthly newsletters announcing items of interest, or answering common questions posed to the RPC.  The AAPM distributes an email burst each month to announce the available of the latest newsletter.

We also continue to maintain the AAPM/RPC registry of brachytherapy seeds meeting the AAPM prerequisite.

General QA activities

The RPC is presently monitoring 1,320 megavoltage therapy sites.  As reported earlier, we reduced the mailing frequency to annually effective January 1, 2003.  We continue to monitor all photon beams with each mailing, and now monitor three electron beams rather than two.  This insures that in most cases, each electron beam is checked no less frequently than biannually.  A recent change is that we now correct our calculations for calibration protocol and calibration medium, for our internal assessment.  In this way, we are able to evaluate if an institution falls within our criteria using the TG-51 protocol to muscle, but falls outside our criteria with the institution's calibration protocol and medium, or vice versa.

As always, we send repeat TLD to institutions whose irradiations fall outside our 5%/5 mm criteria (4% for first time measurements).  We also are following up on non-returned TLD considerably more aggressively than in the recent past.

As of the end of May, 486 institutions have reported that they have converted to the TG-51 protocol;  this is roughly 44% of the U.S. institutions we monitor. The rate of conversion has been steady for the last several years as shown in the graph.

We continue to participate in the Subcommittee on QA of Clinical Trials.

An RPC physicist (Andrea Nelson) recently completed a six month assignment to the MDACC clinic. A second physicist (Francisco Aguirre) began his rotation in July and will finish in December. The program is considerably more effective than before because the RPC physicist fills a clinical physics position in the group supporting the GU, GYN, and breast services.

## Study Group Activities

The RPC continues to send one or more representatives to meetings of every cancer trials study group. We are most heavily involved with NSABP, NCCTG, COMS, GOG, and RTOG. Through coordinated activities with QARC, we are actively involved in reviewing protocols or credentialing activities at several other study groups. In addition, we participate in the QA Committee of ACRIN, the ACR Imaging Network. We participate the Cancer Trials Support Unit (CTSU), to insure that institutions that participate in clinical trials through CTSU are properly monitored and credentialed when necessary.

Radiation Therapy Oncology Group

## Prostate brachytherapy credentialing

We recently conducted a survey of institutions previously credentialed for prostate brachytherapy trials.  This was deemed necessary because the P-0232 trials just opened requires electronic data submission; additional credentialing is required for institutions wishing to participate in this trial. In addition, institutions that experience changes in their treatment planning systems, seed model, or key personnel are expected to be recredentialed. Therefore, of the roughly 60 institutions that were credentialed previously for Protocols 98-05 and P-0019, only 4 are presently credentialed to participate in P-0232.

## RPC anthropomorphic phantoms

- Stereotactic brain phantom

As a result of  ACOSOG informing member institutions that they should obtain our phantom to participate in stereotactic protocols, we believe we may have seen one or two requests for the phantom recently.

2.	Pelvis Phantom

RTOG recently amended the protocol P-0126 to allow IMRT. Four institutions have indicated their interest in becoming credentialed for IMRT using the RPC Pelvic Phantom.  Therefore, three phantoms have been constructed (in addition to the prototype) and we anticipate sending these phantoms in the near future.

- RPC lung phantom

In anticipation of one or more advanced technology protocols for treatment of lung tumors, the RPC is building three additional chest phantoms.  These will be available in the near future for credentialing institutions to participate in these protocols.

- Head and Neck IMRT phantom

The RPC now has a great deal of experience with Head & Neck IMRT Phantoms. Six of these phantoms have been in use; three more are being built by the MD Anderson machine shop in the near future. Beginning with the new phantoms, but extending to the existing phantoms, changes will be made to increase the number of TLD readings and position the TLDs closer to the film. The phantom has been sent to the 43 institutions listed in the Table.  Twenty-one institutions met the evaluation criteria proposed by RPC and recently accepted the RTOG Medical Physics Committee;  agreement between the TLD and the institution's stated dose to the TLDs in the PTV should be within 7%; agreement between the institution's stated dose and the RPC's determination of dose from radiochromic film should be within 4 mm, in the region of high dose gradient between the primary PTV and organ at risk.

Of the 43 institutions that have received the phantom, 5 presently have the phantom to irradiate, and 5 are presently under evaluation. The remaining 12 did not meet the evaluation criteria and are either presently awaiting another phantom, or are re-evaluating their IMRT programs in preparation for another test. At present, 17 institutions are awaiting the availability of a IMRT Head and Neck Phantom, of which 8 are institutions that had previously irradiated the phantom.

| IMRT H & N |
| --- |
| Inst name | Status |
| Akron City Hospital | In Progress |
| Albuquerque Medical Ctr | In Progress |
| CA Cancer Center | In Progress |
| Central Wisconsin Cancer Program | In Progress |
| Cross Cancer Inst | Pass |
| Flower Memorial Hospital | In Progress |
| Fox Chase Med Ctr | In Progress |
| Georgetown Univ Medical Ctr | In Progress |
| Good Samaritan Hospital, Portland | Pass |
| Good Samaritan Medical Center | In Progress |
| Holmes Regional | In Progress |
| Intercommunity Cancer Center | In Progress |
| Iowa Methodist Medical Center | In Progress |
| LDS Hospital | In Progress |
| Mass General Hospital | Pass |
| Mayo Clinic, Rochester | Pass |
| MD Anderson | Pass |
| Medical College of WI | Pass |
| Naval Medical Center Portsmouth | Pass |
| Norfolk General | In Progress |
| Northwest Community Hospital | In Progress |
| Queen's Medical Center, Honolulu, Hawai | Pass |
| Regional Rad Oncology Ctr | In Progress |
| ROC Las Vegas | In Progress |
| Roper Hospital | In Progress |
| Saskatoon | Pass |
| SKCC at Johns Hopkins | Pass |
| St Lukes, Bethlehem, PA | Pass |
| Sutter Cancer Center | In Progress |
| Thomas Jefferson University | Pass |
| Tom Baker Cancer Center | Pass |
| University of Wisconsin | Pass |
| UCSF | Pass |
| University of Floriday | Pass |
| Univ of Michigan Med Ctr | Pass |
| University Hospital of Cincinnati | In Progress |
| University of Albama Medical Ctr | Pass |
| UTMB, Galveston | Pass |
| Virginia Mason Med Ctr | In Progress |
| Washington University | Pass |
| Waukesha Memorial Hospital | Pass |
| Wilkes Barre General Hospital | In Progress |
| William Beaumont Hospital | In Progress |

### Gynecological Oncology Group

The RPC reviews many charts from the GOG and participates extensively in the semi-annual meetings at which clinical reviews are conducted. The GOG is investigating advanced technology radiation therapy, and considering using one or more of our credentialing procedures in the near future. In addition, the GOG has expressed interest in electronic data transfer as departments gradually become filmless, and a preliminary test has been conducted by the RPC to use tools developed by the ATC for this purpose.

The GOG is proposing to open one or more of their protocols internationally.  In particular, they plan to open GOG 191 in Scandinavia. The RPC has been in contact with QA offices in Europe as well as with the NCI to coordinate the monitoring of international institutions.

North Central Cancer Treatment Group

The RPC is participating in the rapid review of cases submitted to N0028, a phase I/II study of escalating doses of radiotherapy for non-small cell lung cancer. NCCTG also has indicted their intention to have participating radiotherapy facilities credentialed for 3D CRT using the benchmark posted on the RPC website.

NEw Projects

Monte Carlo

The RPC is investigating the use of Monte Carlo calculations to enhance the value and range of our database of standard beam data. Several presentations have shown that Monte Carlo can predict accurately many of they dosimetry parameters measured by RPC physicists. We expect that Monte Carlo calculations will supplement our measured data or allow us to reduce the number of measurements we make.

Proton Therapy

As more institutions install proton accelerators, we anticipate the development of NCI-sponsored multi-institutional clinical trials. The RPC is beginning an investigation of dosimetry systems that may be suitable for monitoring participating institutions.

Respiratory Motion

NCI has published recommendations discouraging the use of IMRT in clinical trials for sites that are affected by cardiac or respiratory motion. A number of institutions and manufacturers are investigating gating and tracking techniques to accommodate such motion. The RPC is investigating the design of a phantom to evaluate their techniques. Preliminary discussions have been held with a manufacturer and with representatives of MDACC who are interested in these issues.

## PARTICIPANT FEE:

## Institutions invoiced FY03	1303

No XRT/Canceled/Inactive	91

## Invoiced by RDS	11

Institutions paid	959

### PUBLICATIONS AND ABSTRACTS

Publications Accepted/Published (2001-2003):

- Melia, Sc.M., Michele; Abramson, M.D., David; Albert, M.D., Daniel; Boldt, M.D. Culver; Earle, M.D., John; Hanson, Ph.D., William; Montague, Paul; Moy, Ph.D., Claudia; Schachat, M.D., Andrew; Simpson, M.D., Rand; Straatsma, M.D., Bradley; Vine, M.D., Andrew; and Weingeist, M.D., Ph.D., Thomas: Collaborative Ocular Melanoma Study (COMS) Randomized Trial of I-125 Brachytherapy for Medium Choroidal Melanoma I. Visual Acuity after 3 Years, COMS Report No. 16”, Opthalmology, 108(2):348-366, 2001.

- Kim, C.-H., Reece, W. D., and Cho, S. H.  Computer simulation of radiation exposure in a S/G channel head.  Trans. Am. Nucl. Soc., 84:325-326 , 2001.

- Kirsner, S.M., Prado, K.L., Tailor, R.C., and Bencomo, J.A.: Verification of the accuracy of 3D calculations of breast dose during tangential irradiation: measurements in a breast phantom. Journal of Applied Clinical Medical Physics, Volume 2, Number 3, pp. 149-156, Summer, 2001.

- Tailor R, Hanson W., Calculated absorbed-dose ratios, TG-51/TG-21, for most widely used cylindrical and parallel-plate ion chambers over a range of photon and electron energies, Medical Physics 29:1464-1472, 2002.

- Gifford, K.A., Followill, D.S., Liu, H.H., and Starkschall, G.  Verification of the accuracy of a photon dose-calculation algorithm.  J. Applied Clin. Med. Phys. 3:26-45, 2002.

- Villarruel, S., Ibbott, G.S., and Lai-Fook, S.J.: Effect of concentration and hydration on restriction of albumin by lung interstitium.  Microvascular Research 63, 27-40, 2002.

- Cadman, P., Bassalow, R., Sidhu, N.P.S., Ibbott, G., Nelson, A.: Dosimetric considerations for validation of a sequential IMRT process with a commercial treatment planning system.  Physics in Medicine and Biology Vol. 47, 3001-3010, 2002.

- Aguirre, J.F., Tailor, R., Ibbott, G., Stovall, M. Hanson, W.  TLD as a tool for remote verification of output for radiotherapy beams: 25 years of experience.  Accepted by International Atomic Energy Agency, November, 2002.

- Ibbott, G., Beach, M., Maryanski, M.  An anthropomorphic head phantom with a BANG® polymer gel insert for dosimetric evaluation of IMRT treatment delivery. Accepted by International Atomic Energy Agency, November, 2002.

- Ibbott, G., Nelson, A., Followill, D., Balter, P., Hanson, W.  An anthropomorphic head and neck phantom for evaluation of intensity modulated radiation therapy. Accepted by International Atomic Energy Agency, November, 2002.

- Izewska, J., Svensson, H., Ibbott, G.  Worldwide QA networks for radiotherapy dosimetry. Accepted by International Atomic Energy Agency, November, 2002.

- Tailor R, Hanson W, and Ibbott G, TG-51 Experience from 150 institutions, common errors, and helpful hints, Journal of Applied Clinical Medical Physics, 4:102-111, 2003.

- Urie, M., FitzGerald, T.J., Followill, D., Laurie, F., Marcus, R., Michalski, J.  Current calibration, treatment, and treatment planning techniques among institutions participating in the Children’s Oncology Group. Int. J. of Radiat. Oncol. Biol. Phys. 1:245-260, 2003.

- Followill D.S., Hanson, W.F., Ibbott, G.S., Eglezopoulos, L.R., and Chui, C.S.  Differences in electron beam dosimetry using two commercial ionization chambers and the TG-21 protocol: another reason to switch to TG-51.  J. Applied Clin. Med. Phys. 4: 124-131, 2003.

- Nag, S., Quivey, J.M., Earle, J.D., Followill, D.S., Fontanesi, J., and Finger, P. The American Brachytherapy Society Recommendations for Brachytherapy of Uveal Melanomas, Int. J. of Radiat. Oncol. Biol. Phys. 56:544-555, 2003.

- Krintz, A.L., Hanson, W.F., Ibbott, G.I. and Followill, D.S., A Reanalysis of the Collaborative Ocular Melanoma Study Medium Tumor Trial Eye Plaque Dosimetry, Int. J. of Radiat. Oncol. Biol. Phys. 56:889-898, 2003.

- Diener-West, M., Albert, D. M., Frazier Byrne, Sl, Davidorf, F. H. Followill, D. S., Green, R.L., Hawkins, B.S., Kaiser, P.K., Robertson, D.M., and Straatsma, B.R., Comparison of Clinical, Echographic and Histopathologic Measurements from Eyes with Medium-Sized Choroidal Melanoma: in the Collaborative Ocular Melanoma Study.  COMS Report No. 21, The Collaborative Ocular Melanoma Study Group, Archives of Ophthalmology 121: 2003.

- Tailor, R.C., Followill, D.S., Hernandez, N., Ibbott, G.S., and Hanson, W.F., “Predictability of electron cone ratios with respect to linac make and model,” Journal of Applied Clinical Medical Physics, Vol. 4 (2), pp. 172-178, February 2003.

- Followill, D.S., Stovall, M.S., Kry, S.F., and Ibbott, G.S., Neutron source strength measurements for Varian, Siemens, Elekta, and General Electric linear accelerators. Accepted by the J. Applied Clin. Med. Phys., 2003.

- Bencomo, J.A., Tello, V.M., Chu, C., Cho, S.H., and Ibbott, G.S.: Anthropomorphic Breast Phantoms for Quality Assurance and Dose Verification. Journal of Applied Clinical Medical Physics (In review, 2003).

Abstracts

- Followill, D. and Stovall, M: Tinea Capitis: Uncertainties in radiation dose estimates.  Radiation Research, Vol. 154, p.720-721, 2000.

- Balter P, Lowenstein J, and Hanson W: Electron Calibrations: Parallel Plate Chambers vs. Cylindrical Chambers Using TG-51. Medical Physics, Vol. 28: 1214, 2001.

- Radford, D.A., Followill, D.S., and Hanson, W.F.: A Standard Method of Quality Assurance for Intensity Modulated Radiation Therapy of the Prostate. Medical Physics 28 (6): 1211, 2001.

- Followill, D.S., Hernandez, N., and Hanson, W.F.: Standard Wedge and Tray Transmission Values for Varian, Seimens, Elekta/Philips Accelerators; A Quality Assurance Tool. Medical Physics 28 (6): 1211, 2001.

- Krintz, A., Followill, D., Melia, M., and Hanson W.F.: A Reanalysis of the Collaborative Ocular Melanoma Study Medium Tumor Trial Eye Plaque Dosimetry. Medical Physics 28 (6): 1192, 2001.

- Tailor, R.C. and Hanson, W.F.: Estimated Absorbed - Dose Ratios "TG51/TG21" for Most Commonly Used Cylindrical and Parallel-Plate Ion Chambers over a Range of Photon and Electron Energies. Medical Physics 28 (6): 1192, 2001.

- BenComo, J.A., Hernandez, N., and Hanson, W.F.: Problems and Shortcomings of the RPC Remote Monitoring Program of Institutions Dosimetry Data. Medical Physics 28 (6): 1212, 2001

- Hanson, W.F., Martin, B., Kuske, R., Arthur, D., Rabinovitch, R., White, J., Wilenzick, R. Harris, I., Tailor, R., and Davis, D.: Dose Specification and Quality Assurance of RTOG Protocol 95-17, A Cooperative Group Study of 192Ir Breast Implants as Sole Therapy. Medical Physics 28 (6): 1297, 2001.

- Lowenstein J, Roll J, Hanson W, Davis D, Lanciano R, Calkins A, Petereit, D, Varia M, and Ibbott, G, Radiotherapy Quality Assurance of Gynecologic Oncology Group (GOG) Protocol 165, A Cooperative Group Study of Carcinoma of the Cervix, Int. J. Rad. Onc., Bio., Phys., 54:2, 283, 2002.

- Cho, S. H. and Ibbott, G.S.  Reference Dosimetry Data for the Siemens Primus Linear Accelerator: Preliminary Results for Depth Dose and Output Factor, Medical Physics 29:1944, 2002.

- Krintz, A., Hanson, W., Ibbott, G., and Followill, D., Verification of PLAQUE SIMULATOR dose distributions using radiochromic film.  Medical Physics 29: 1220-1221, 2002.

- Kry, S., Salehpour, M., Followill, D., Stovall, M., and Rosen, I., Secondary Dose from IMRT treatments.  Medical Physics 29:1348, 2002.

- Krintz, A., Hanson, W., Ibbott, G., and Followill, D., A reanalysis of the Collaborative Ocular Melanoma Study medium tumor trial eye plaque dosimetry. Int. J. Radiat. Oncol. Biol. Phys. 54: 329-330, 2002.

- Duzenli, C., Baldock, C., Ibbott, G., Schreiner, L.J., Jordan, K., Oldham, M., Short, K., CCPM/COMP Symposium on Gel Dosimetry, AAPM 44th Annual Meeting, Montreal, Quebec, Canada, 2002.

- Nath, R., Rivard, M., Coursey, B., DeWerd, Hanson, W., Huq, M., Ibbott, G., Williamson, J., Status of the American Association of Physicists in Medicine Radiation Therapy Committee’s Subcommittee on Low-Energy Interstitial Brachytherapy Source Dosimetry: Procedure for the Development of Consensus Single-Source Dose-Distribution. AAPM 44th Annual Meeting, Montreal, Quebec, Canada, 2002.

- Beach, M.L., Ibbott, G.S., and Maryanski, M., Implementation of a Polymer Gel Dosimetry Insert for an Anthropomorphic Head Phantom used to Evaluate Head and Neck Intensity-Modulated Radiation Therapy, AAPM 2002.

- Cho, S.H. and Ibbott, G.S., Reference Photon Dosimetry Data for the Siemens Primus Linear Accelerator: Preliminary Results for Depth Dose and Output Factor, AAPM 2002.

- Aguirre, J.F., Tailor, R., Ibbott, G., Stovall, M. Hanson, W.  TLD as a tool for remote verification of output for radiotherapy beams: 25 years of experience.  Accepted by International Atomic Energy Agency, November, 2002.

- Ibbott, G., Beach, M., Maryanski, M.  An anthropomorphic head phantom with a BANG® polymer gel insert for dosimetric evaluation of IMRT treatment delivery. Accepted by International Atomic Energy Agency, November, 2002.

- Ibbott, G., Nelson, A., Followill, D., Balter, P., Hanson, W.  An anthropomorphic head and neck phantom for evaluation of intensity modulated radiation therapy. Accepted by International Atomic Energy Agency, November, 2002.

- Izewska, J., Svensson, H., Ibbott, G.  Worldwide QA networks for radiotherapy dosimetry. Accepted by International Atomic Energy Agency, November, 2002.

- Nelson, A., Balter, P., Hanson, W., Ibbott, G., Experience with an IMRT Head and Neck QA Phantom.  Medical Physics, 29:1366, 2002.

- Bencomo, Jose A., and Ibbott, Geoffrey S., Quality Assurance Measurements of Beam parameters of a Linear Accelerator using MOSFET Dosimeters: A feasibility Study. Medical Physics, 29, (6): 1201, 2002.

- Lowenstein J, Roll J, Ibbott G: To Plan or Not to Plan: Multiple High Dose Rate (HDR) Brachytherapy Insertations, Medical Physics, 30:1464, 2003.

- Bayouth, J., Followill, D., Nelson, A., Ibbott, G., and Morrill, S., Dosimetric Effect of a Hip Prosthesis when Delivering Intensity Modulated Radation Therapy (IMRT):  A Phantom Study, poster at ESTRO Annual meeting September 2003.

- Kry, S., Salehpour, M., Followill, D., Stovall, M., Rosen, I. Risk Assessment of Secondary Malignancies from IMRT Treatments, Medical Physics, 30 (6): 1130, 2003.

- Heard, M., Ibbott, G., Followill, D.  Characterizing Dose Distributions of Brachytherapy Sources Using Normoxic Gel (WIP), Medical Physics, 30 (6): 1333, 2003.

- Esteban, J., Ibbott, G., Hanson, W., Horton, J., White, A., Cho, S.  Energy Dependence of a TLD System for Characterizing Low Energy Brachytherapy Sources (WIP), Medical Physics, 30 (6): 1349, 2003.

- Bencomo, J., Ibbott, G., Lee, S., Lii, M., Borges, J.  Evaluation of Two Linear Ionization-Chamber-Arrays as Data Acquisition Systems for Quality Assurance of Photon and Electron Beams, Medical Physics, 30 (6): 1350, 2003.

- Beach, M., Ibbott, G., Followill, D., Hanson, W., Bloch, C., Jackson, E., Tucker, S.  Implementation of a Polymer Gel Dosimetry Insert for An Anthropomorphic Phantom Used to Evaluate Head and Neck Intensity-Modulated Radiation Therapy, Medical Physics, 30 (6): 1396, 2003.

- Rivard, M.J., Coursey, B.M., DeWerd, L.A., Hanson, W.F., Huq, M.S., Ibbott, G.S., Nath, R., Williamson, J.F.  Update of AAPM Task Group No. 43 Report - A Revised Protocol for Brachytherapy Dose Calculations, Medical Physics, 30 (6): 1431, 2003.

- Rogers, D., DeWerd, L., Ibbott, G., Huq, M.  Changes in Co-60 Air-Kerma Standards: The Rationale for Change and the Impact On Clinical Practice, Medical Physics, 30 (6): 1442, 2003.

- Vassiliev, O., Liu, H.H., Dong, L, Antolak, J., Cho, S., Mohan, R.  A Monte-Carlo-Based System for Dose Calculation and Dosimetry QA of Photon IMRT (WIP), Medical Physics, 30 (6): 1453, 2003.

- Cho, S.H., Kim, C.H., Ibbott, G.S.  Differences in the Results of MCNP Simulation for Low Energy Photon Sources  May Be Due to the Choice of Photon Cross Section Libraries, Medical Physics, 30 (6): 1464, 2003.

- Tailor, R., Esteban, J., Ibbott, G.S.  TLD Measurements in Liquid Water of Dosimetry Characteristics of a New 125I Seed, Medical Physics, 30 (6): 1470, 2003.

- Borges, J.A., Bencomo, J., Ibbott, G.S.  A 3 Dimensional Gel Dosimetry Lung Equivalent (WIP), Medical Physics, 30 (6): 1478, 2003.

- Tailor, R., Followill, D.S., Hernandez, N., Zhu, T.S., Ibbott, G.S.  Quality-Assurance Check of Collimator and Phantom-Scatter Factors, Medical Physics, 30 (6): 1483, 2003.

- Fisher, G., Followill, D., Tolani, N., Ibbott, G.  The Accuracy of 3-D Inhomogeneity Photon Algorithms in Commercial Treatment Planning Systems Using a Heterogeneous Lung Phantom (WIP), Medical Physics, 30 (6): 1507, 2003.

- Tailor, R., Hanson, W., Wells, N., Ibbott, G.  Are Photon and Electron Beam Calibrations More Consistent with TG-51 Than with TG-21?, Medical Physics, 30 (6): 1350, 2003.

Invited Articles

- Cho, S. H.  Dosimetry of brachytherapy sources: Review of the AAPM TG-43 formalism, Proceedings of the 3rd Korea-Japan Joint Meeting on Medical Physics and the 2nd Asia Oceania Congress of Medical Physics 2002, ISSN 1589-6691, September 2002.

- Greven, K.M., Levenback, C., Chao, C.K.S., Delaney, T, Del Priore, G., Eifel, P., Erickson, B.A., Followill, D., Gaffney, D., Garcia, M., Gerszten, K., Grigsby, P., Henderson, R., Hricak, H., Hsu, J., Jhingrin, A., Kaye, A., Kudelka, A., Lukka, H., Mutch, D., Nag, S., Rotman, M., Shefter, T., Smith, W., Stehman, F., Souhami, L., Wenzel, L., Winter, K.A., and Wolfson, A. “Gynecologic Cancer Working Group,” Int. J., Radiation Oncology Biol. Phys, Vol. 51 No. 3 Supplement 2, 58-59, 2001.

- Michalski, J., Purdy, J.A., Gaspar, L., Souhami, L., Ballow, M., Bradley, J., Chao, C.K.S., Crane,

C., Eisbruch, A., Followill, D., Forster, K., Fowler, J., Gillin, M.T., Graham, M.L. Harms, W.B., Huq, S., Kline, R.W., Langer, M., Mackie, T.R., Mukherji, S., Podgorsak, E.B., Roach, M., Ryu, J., Sandler, H., Schultz, C.J., Schell, M., Verhey, L.J. Vicini, F., and Winter, K.A., “Image-Guided radiation Therapy Committee”, Int. J., Radiation Oncology Biol. Phys, Vol. 51 No. 3 Supplement 2, 60-65, 2001.

Letters to the Editor/Newsletters

- Ibbott, G.S. and Nath, R.: “Dose-rate constant for Imagyn 125I brachytherapy source.  Medical Physics 28(4), April 2001.

- Ibbott, G.S., Meigooni, A.S., and Gearheart, D.M.: Monte Carlo determination of dose rate

constant,” Medical Physics 29, 1637-1638, 2002.

- Rivard, M.J., Coursey, B.M., DeWerd, L.A., Hanson, W.F., Huq, M.S., Ibbott, G.S., Nath, R.,

Williamson, J.F.: Comment on “Let’s abandon geometry factors other than that of a point source in brachytherapy dosimetry,” Medical Physics 29, 1917-1918, 2002.

- Followill, D.S., Lowenstein, J.R., and Ibbott, G, Quality Assurance: It’s Here to Stay, Newsletter American Association of Physicists in Medicine, Vol. 28, No. 1, January/February 2003.

- Tailor R, Hanson W, Ibbott G. Response to Comment on Calculated absorbed-dose ratios, TG51/TG21, for most widely used cylindrical and parallel-plate ion chambers over a range of photon and electron energies. Letter to Editor, Medical Physics, 30:478-480,2003.

- Olch, A., Followill, D., Thomadsen, B., Wichman, B., Rosemark, P., and Hunt, M., In regard to Urie et al.: Current calibration, treatment, and treatment planning techniques among institutions participating in the children'’ oncology group. Int. J., Radiation Oncology Biol. Phys, Vol. 56 No. 4:1209-1210, 2003.

Manuals, Teaching Aids, Other Publications

- Digital Imaging Technology Applications in Mammography in CD of proceedings of the 2nd meeting              of ALFIM in Caracas Venezuela, October 2001.

- Bencomo, J.A., Hernandez, N., and Hanson, W.F. Problems and Shortcomings of the RPC Remote Monitoring Program of Institutions Dosimetry Data. AAPM 43rd Annual Meeting, Salt Lake City, Utah, July 22-26, 2001.

- Bencomo, Jose A., and Ibbott, Geoffrey S., Quality Assurance Measurements of Beam parameters of a Linear Accelerator using MOSFET Dosimeters: A feasibility Study. AAPM 44th Annual Meeting, Montreal, Ontario, Ca, June, 13-18, 2002.

- Lowenstein, J., Educational video on how to utilize the TG-51 Calibration Protocol, 2001.

- Ibbott, G., Manual on Radiation Protection in Hospitals and General Practice, ed. by Gerald P. Hanson, to be published by WHO, 2001.

- Aguirre, F., Physical aspects of quality assurance in radiotherapy recommendation from a panel of experts from the IAEA within the activation of the Regional Program ARCAL XXX (to be published).

7.	Aguirre, F. and Ibbott, G., Lecture materials for "Prevention of Accidental Exposure", IAEA Division of Radiation and Waste Safety, 2003.

Book Chapters

- Ibbott, G., What every medical physicist should how about the JCAHO standard. In Accreditation Programs and the Medical Physicist.  ed. by R.L. Dixon, P.B. Butler, and W.T. Sobol, AAPM Monograph Series, Medical Physics Publishing, 2001.

- Ibbott, G., Professional rols in VBT. In Intravascular Brachytherapy/Fluoroscopically Guided Interventions.  ed. by Stephen Balter, Rosanna C. Chan, and Thomas B. Shope, Jr., AAPM Medical Physics Monograph No. 28, 2002.

- BenComo, José A., Introduction to Radiation Protection for Medical Physicists, Text Book CD-ROM (under development)

Books (edited and written)

- Radiation Therapy Physics, 2nd edition, W.R. Hendee and G.S. Ibbott, Mosby Year-Book Publishers, Philadelphia, 1996. (under contract to prepare 3rd edition to be published by John Wiley, Inc.)

- Principles of Digital Mammography 1st Ed. (University Autonomous of Mexico City, Mex. (under Development), 2002.

## PRESENTATIONS

INTERNATIONAL ACTIVITIES

### Geoffrey Ibbott attended the International Electrotechnical Commission Meeting in Crawley, England, May 11 – 14, 2003 – (all expenses borne by the American College of Radiology)

Geoffrey Ibbott did a presentation of the "RPC Phantoms" to the Computerized Imaging Reference Systems, Inc. (CIRS), Norfolk, Virginia, March 20 – 21, 2003.

### VISITS TO INSTITUTIONS  (March 1 - July 31, 2003)

- Sang Cho performed radiological physics measurements and reviewed patient dosimetry at the University of Maryland, Baltimore, Maryland, March 9 – 15, 2003.

- Jessica Lowenstein performed radiological physics measurements and reviewed patient dosimetry at the University of Wisconsin Medical Center, Madison, Wisconsin, March 10 – 13, 2003.

- David Followill performed radiological physics measurements and reviewed patient dosimetry at MD Anderson Cancer Center, Houston, TX, April 7 - May 20, 2003.

- Ramesh Tailor performed radiological physics measurements and reviewed patient dosimetry at Via Christi and St. Francis Medical Centers, Wichita, Kansas, May 16 – 23, 2003.

- Francisco Aguirre performed radiological physics measurements and reviewed patient dosimetry at Kaplan Cancer Center, West Palm Beach, FL, June 3 – 6, 2003.

- Jessica Lowenstein performed radiological physics measurements and reviewed patient dosimetry at the Cleveland Clinic, Cleveland, Ohio, June 8 – 11, 2003.

- Paola Alvarez performed radiological physics measurements and reviewed patient dosimetry at the Cleveland Clinic, Cleveland, Ohio, June 9 – 11, 2003.

- Geoffrey Ibbott performed radiological physics measurements and reviewed patient dosimetry at Swedish Medical Center, Englewood, Colorado, July 18 - 23, 2003

- Andrea Nelson performed radiological physics measurements and reviewed patient dosimetry at Cross Cancer Center, Edmonton, Alberta, Canada, July 28 - August 1, 2003.

- Jessica Lowenstein performed radiological physics measurements and reviewed patient dosimetry at Cross Cancer Center, Edmonton, Alberta, Canada, July 28 - August 1, 2003.

### MEETINGS ATTENDED (March 1 - July 31, 2003)

- David Followill attended the Childrens Oncology Group Meeting in Atlanta, Georgia, March 26 – 29, 2003.

- Andrea Nelson attended the Southwestern Chapter Meeting of the American Association of Physicists in Medicine in Houston, Texas, March 28 – 29, 2003.

- Geoffrey Ibbott attended the Southwest Oncology Group Meeting in San Diego, California, April 10 - 12, 2003.

- Francisco Aguirre attended the Southwest Oncology Group Meeting in San Diego, California, April 10 - 13, 2003.

- David Followill attended the Advanced Technology Consortium Meeting in Chicago, Illinois, April 15 – 16, 2003.

- Irene Harris attended the North Central Cancer Treatment Group Meeting in Rochester, Minnesota, April 23 - 24, 2003.

- Geoffrey Ibbott attended the North Central Cancer Treatment Group Meeting in Rochester, Minnesota, April 23 - 24, 2003.

- Jose Bencomo attended the Advanced Technology Consortium / Digital Imaging in Communications and Medicine Meeting in St. Louis, Missouri, May 2 – 3, 2003.

- Huy Duong attended the Advanced Technology Consortium / Digital Imaging in Communications and Medicine Meeting in St. Louis, Missouri, May 2 – 3, 2003.

- Sang Cho attended the Eastern Cooperative Oncology Group Meeting in New Orleans, Louisiana, June 5 - 6, 2003.

- Geoffrey Ibbott attended the American College of Surgeons Oncology Group Meeting in Montreal, Quebec, Canada, June 5 – 7, 2003.

- Geoffrey Ibbott attended the Radiation Therapy Oncology Group Meeting in Montreal, Quebec, Canada, June 25 – 29, 2003.

- David Followill attended the Radiation Therapy Oncology Group Meeting in Montreal, Quebec, Canada, June 25 – 29, 2003.

- Joye Roll attended the Radiation Therapy Oncology Group Meeting in Montreal, Quebec, Canada, June 25 – 29, 2003.

- Ramesh Tailor attended the National Surgical Adjuvant Breast and Bowel Project Meeting in Orlando, Florida, June 26 – 30, 2003.

- Cindy Davis attended the National Surgical Adjuvant Breast and Bowel Project Meeting in Orlando, Florida, June 26 – 30, 2003.

- Geoffrey Ibbott attended the Gynecological Oncology Group Meeting in Reno, Nevada, July 24 - 26, 2003.

- Jessica Lowenstein attended the Gynecological Oncology Group Meeting in Reno, Nevada, July 24 - 26, 2003.

- Joye Roll attended the Gynecological Oncology Group Meeting in Reno, Nevada, July 24 - 26, 2003.

### STATUS SUMMARY

Dosimetry review visits were made to 9 institutions, at which 80 beams† were reviewed, 3,197 TLD reports were evaluated, and 644  patients' charts were reviewed. There are five institutions that are a special high priority for a dosimetry review visit (based on TLD, chart check, etc).

A summary of cooperative groups and institutions’ participation in them are included in the attached table.

†The review of one photon energy is considered a "beam"; however, review of a cadre of electrons on one machine is considered only one "beam".  Measurement of TBI and other special procedures are also considered one "beam".

Respectfully submitted,

Geoffrey S. Ibbott, Ph.D.

Director, Radiological Physics Center

EXPENDITURES OF THE RADIOLOGICAL PHYSICS CENTER

(RPC Grant, the COMS Subcontract, and the High Technology Subcontract)

March 1, 2003 - July 31, 2003

____________________________________________________________________________

## PERSONNEL (salaries, fringe benefits):	$957,022.48

10 Physicists, 1 Supervisor of Quality Assurance Dosimetry Services, 4 Dosimetrists, 1 Information Manager, 1 Database Administrator, 1 Programmer Analyst I, 1 Radiological Physics Supervisor, 6 Radiological Physics Technicians, 1 Department Administrator, 1 Sr. Administrative Assistant, 1 Departmental Buyer, 3 Secretaries, and 4 Graduate Research Assistants.

______________________________________________________________________________

### TRAVEL

### Visits

| Sang Cho – University of Maryland | $1667.05 |
| --- | --- |
| Jessica Lowenstein – University of Wisconsin Medical Center | $928.41 |
| David Followill - MD Anderson Cancer Center | $0.00 |
| Ramesh Tailor – Via Christi and St. Francis Medical Centers | $1472.53 |
| Francisco Aguirre – Kaplan Cancer Center | $1036.02 |
| Jessica Lowenstein – Cleveland Clinic | $900.45 |
| Paola Alvarez – Cleveland Clinic | $758.06 |
| Geoffrey Ibbott – Swedish Medical Center | $770.24 |
| Andrea Nelson – Cross Cancer Center | $2165.00 |
| Jessica Lowenstein – Cross Cancer Center | $2165.00 |
| TOTAL: | $11862.76 |

Meetings:

| David Followill – COG | $633.56 |
| --- | --- |
| Andrea Nelson – SWAAPM | $80.00 |
| Geoffrey Ibbott – SWOG | $773.56 |
| Francisco Aguirre – SWOG | $1025.80 |
| David Followill – ATC | $646.97 |
| Irene Harris – NCCTG | $589.88 |
| Geoffrey Ibbott – NCCTG | $751.55 |
| Jose Bencomo – ATC/DICOM | $508.56 |
| Huy Duong – ATC/DICOM | $450.76 |
| Sang Cho – ECOG | $521.70 |
| Geoffrey Ibbott – ACOSOG | $987.95 |
| Geoffrey Ibbott – RTOG | $1255.33 |
| David Followill – RTOG | $1477.61 |
| Joye Roll – RTOG | $1241.47 |
| Ramesh Tailor – NSABP | $955.00 |
| Cindy Davis – NSABP | $1037.53 |
| Geoffrey Ibbott – GOG | $781.85 |
| Jessica Lowenstein – GOG | $767.48 |
| Joye Roll – GOG | $794.35 |
| TOTAL: | $15280.91 |

| CONSULTANTS Fiscal 2003 this funds contains $12,500 James Balter (Houston) Arthur Olch (Houston) Howard Omols (Houston) Timothy Zhu (Houston) Saiful Huq (Houston) Zuofeng Li (Houston) Marc Sontag (Houston) Jeffrey Williamson (Houston) Dennis Duggan (Houston) Jatinder Palta (Houston) Daniel Low (Houston) Gary Ezzell (Houston) Michael Herman (Houston) Rasika Rajapakshe (Houston) Chang-Ming Ma (Houston) TOTAL: | $ 703.26 779.77 921.28 886.73 630.89 584.04 744.44 800.73 677.85 230.55 517.06 575.05 1092.10 1060.95 858.86 11,063.56 |
| --- | --- |
| SUPPLIES: | $ 44,028.36 |
| Office supplies, laboratory and record keeping, TLD, |  |
| TLD supplies, software, etc. |  |
| OTHER EXPENSE: | $ 14,016.46 |
| Postage, telephone, reprints, copying, computer fees, |  |
| equipment repair, etc. |  |
| SPACE RENTAL: | $ 80,976.25 |
| Total Expenditures March 1, 2003 - July 31, 2003 | $1,107,107.11 |
| Indirect costs @ 25% | 276,776.78 |
| TOTAL | $1,383,883.89 |

RPC Report to RTC September 2002

| Clinical Study Groups | Office Reviewing Patient Records | Special Projects |
| --- | --- | --- |
| American College of Radiology Imaging Network ACRIN | N/A | Participate in the development of guidelines for quality assurance of institution participating in ACRIN CT Dose Measurements |
| American College of Surgeons Oncology Group ACOSOG | QARC/RPC | TRUS Prostate Approval Collaboration RPC Institution List Stereotactic Phantom |
| Cancer and Acute Leukemia Group B CALGB | QARC | TRUS Prostate Approval Collaboration |
| Children’s Oncology Group COG (POG, CCG, WILMS, RHABDO) | QARC | 3D Benchmark Case IMRT Benchmark Case |
| Eastern Cooperative Oncology Group ECOG | QARC |  |
| Gynecologic Oncology Group GOG | RPC | Radiotherapy manual HDR Pre-approval Electronic Transfer of Patient Records Image Based Treatment Planning IMRT Guidelines |
| National Surgical Adjuvant Breast and Bowel Project NSABP | RPC | IMRT Guidelines Breast phantom Pre-approval discussion |
| North Central Cancer Treatment Group NCCTG | RPC | Rapid Review of Lung Study TRUS guided Prostate Pre-approval 3D CRT credentialing |
| Radiation Therapy Oncology Group RTOG | RTOG/RPC | IMRT Phantom /pre-approval Prostate Phantom Prostate Implant Pre-approval LDR/HDR Stereotactic Phantom Lung Phantom Cervix HDR Pre-approval Patient Calculation Program |
| Southwest Oncology Group SWOG | QARC |  |
| Clinical Trial Support Unit CTSU | QARC, RPC, RTOG | RPC Institution List RTF Numbers TLD Monitoring |
| Collaborative Ocular Melanoma Study COMS | RPC | Cataracts follow-up Visual acuity dose response |