## Slide 1: Investigation of a Commercial OSLD System for CT Dosimetry
S Scarboro1,2, D Cody1,2, D Followill1,2, P Alvarez1,
M McNitt-Gray3, D Zhang3, L Court1,2, S Kry1,2*
1UT MD Anderson Cancer Center, Houston, TX
2UT Health Science Center Graduate School of Biomedical Sciences, Houston, TX
3UCLA School of Medicine, Los Angeles, CA
Tuesday, July 31, 2012 AAPM Annual Meeting
Charlotte, NC

## Slide 2: Motivation
CT dosimetry is an area of increasing interest
CT Dose Index (CTDI) is standard approach for dose quantification in CT
  Not good metric for patient dose
Desire to improve patient dose assessment
Optically Stimulated Luminescence Dosimetry (OSLD)
  High precision, cost effective, doesn’t perturb image
  Common dosimeter in therapy environment
  Unknowns/Issues when applied to CT environment
    Limited characterisation and calibration methods

## Slide 3: Study design
OSLD
  Al2O3:C based chip – “nanoDot”
  Commercially available - Landauer, Inc
Read with MicroStar
  High intensity beam
In this study
  Full characterization of nanoDot
  for CT
    Focus on energy response
  Calibration protocols

## Slide 4: OSLD Dosimetry
OSLD Measurement
(Signal, M)
Vendor Calibration
kd
CT Calibration
kd, kE, kθ
Therapy Calibration
kd, kG, kE, kθ

## Slide 5: Characterizing the nanoDot
 |  | Vendor Calibration | CT Calibration | Therapy Calibration
Calibration Coefficient | CD |  |  | 
Signal Depletion | kd |  |  | 
Signal Fading | kF |  |  | 
Signal Linearity | kL |  |  | 
Irradiation Geometry | kG |  |  | 
Angular Dependence | kθ |  |  | 
Energy Dependence | kE |  |  | 
Project goal: Fill in table
  Calibration Protocol determines CD
  Additional correction factors dependent on calibration protocol used + measurement condition

## Slide 6: Calibration: CD
Vendor Calibration
  Pre-irradiated dosimeter (80kVp beam)
  Constant Energy Correction Factor = 1.19
CT Free-In-Air Calibration
  Irradiate ion chamber (Dose) and OSLD (signal) identically in air in CT
  Corrections to be determined
Therapy Calibration
  Irradiate ion chamber (Dose) and OSLD (signal) identically in MV beam
  Corrections to be determined

## Slide 7: Characterizing the nanoDot
 |  | Vendor Calibration | CT Calibration | Therapy Calibration
Calibration Coefficient | CD | Determine as described…. |  | 
Signal Depletion | kd | ~1.02 | ~1.02 | ~1.02
Signal Fading | kF | -- | 1.00 | 1.00
Signal Linearity | kL | -- | 1.00 | 1.00
Irradiation Geometry | kG | -- | 1.00 | 1.03
Angular Dependence | kθ | -- | 1.00 | 1.00
Energy Dependence | kE | 1.19 |  | 
Minimal corrections required for CT applications

## Slide 8: Energy Correction - kE
Energy Correction represented largest correction for each calibration approach
Consider changes in energy with
  kVp, phantom size, measurement position, scan extent
kE determined two ways
  Theoretical Approach – Burlin Cavity Theory + Monte Carlo Simulated Spectra
  Measurement Approach – Ion Chamber + OSLD

## Slide 9: Calculated Values of kE CT Free-In-Air Protocol
Chart: Mean Spectral Energy (keV)
| 45.350774393992737 | 43.804472368369773 | 46.777948080256962 | 45.884931429865922 | 48.606300418073573 | 47.889936967510323
80 kVp | 0.81661 | 0.80219 | 0.83208 | 0.82241 | 0.85386 | 0.84503
120 kVp | 0.91184 | 0.87458 | 0.92566 | 0.99674 | 0.96585 | 0.96911 | 1.0453 | 1.0174 | 1.0198
140 kVp | 0.95678 | 0.9091 | 0.96756 | 0.92397 | 1.0819 | 1.0374 | 1.0267 | 0.9909 | 1.1458 | 1.1049 | 1.0888 | 1.0559
Measured kE for 11 different scans agreed with calculated values within 5% on average

## Slide 10: Energy Correction (kE)
kE varies with
  kVp, location in phantom, size of phantom, scan extent
kE is within 2-3% based only on kVp and position of measurement
    CT Free-In-Air Protocol

## Slide 11: Characterizing the nanoDot
 |  | Vendor Calibration | CT Calibration | Therapy Calibration
Calibration Coefficient | CD | Determine as described…. |  | 
Signal Depletion | kd | ~1.02 | ~1.02 | ~1.02
Signal Fading | kF | -- | 1.00 | 1.00
Signal Linearity | kL | -- | 1.00 | 1.00
Irradiation Geometry | kG | -- | 1.00 | 1.03
Angular Dependence | kθ | -- | 1.00 | 1.00
Energy Dependence | kE | 1.19 | table lookup 0.81-1.10 | table lookup 0.29-0.39

## Slide 12: Dose Accuracy
Chart: Mean Photon Energy (keV)
| 45.35 | 47.89 | 54.690000000000012 | 59.53 | 55.31 | 60.720000000000013 | 51.660000000000011 | 58.65 | 58.05 | 64.849999999999994 | 63.190000000000012
Ion Chamber | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
Therapy Calibration | 0.95775 | 0.97867 | 0.97323 | 1.009 | 0.94938 | 1.0554 | 0.94555 | 0.94145 | 0.94964 | 1.069 | 1.0414
CT Calibration | 0.9684 | 0.99156 | 0.95969 | 1.0023 | 0.95675 | 1.0693 | 0.95983 | 0.95055 | 0.94113 | 1.0684 | 1.0388
Vendor Calibration (standard) | 0.95577 | 0.9441 | 0.85351 | 0.81115 | 0.83259 | 0.84843 | 0.88273 | 0.80135 | 0.78522 | 0.79324 | 0.79313

## Slide 13: Summary and Conclusion
Vendor calibration showed worse agreement for higher scan energies
  >20% lower dose predicted for 140kVp scans
Can achieve good accuracy with OSLD
 | Average Disagreement with Ion Chamber
CT Free-In-Air Calibration | 4.1%
Therapy Calibration | 4.4%
Vendor Calibration | 15.5%

## Slide 14: Thank you!