## IROC Houston Head and Neck Phantom

Guidelines for Planning and Irradiating the IROC Houston IMRT Phantom.

Revised Dec 2007

Credentialing for this protocol requires three steps: (1) submission of the Facility Questionaire with supporting documentation to the Image Guided Therapy Center (ITC, http://itc.wustl.edu), (2) a successful dry run test, and (3) completion of the phantom treatment experiment. The purpose of steps (2) and (3) is to confirm that the IMRT dose distribution planned by each institution can be delivered by that institution, and correctly submitted to the ITC.

The RTOG is requesting that each institution keep the phantom for a period of time less than 2 weeks.  During this two-week period, the institution will image, plan, and treat the phantom and return it to the IROC Houston.  Thank you for your cooperation.

This phantom has been designed and constructed by the IROC Houston QA Center.  The IROC Houston phantom contains TLD at 4 locations and perpendicular sheets of film.  Two TLD are located in the primary PTV, and one each in the secondary PTV and in the organ at risk (spinal cord).  The guidelines for dose delivery were written by Mike Gillin.

If you have any questions, please contact the appropriate person.

IROC Houston	Geoff Ibbott	(713) 745-8989	gibbott@mdanderson.org

IROC Houston	Andrea Molineu	(713) 745-8989	amolineu@mdanderson.org

MD Anderson	Mike Gillin	(713) 563-2507	mgillin@mdanderson.org

ITC		Jim Purdy 	(314) 362-2639	purdy@castor.wustl.edu

ITC		Bill Straube	(314) 362-9762	itc@castor.wustl.edu

DOSIMETRY INFORMATION TO BE SUBMITTED:

The following information is to be submitted to the IROC Houston (include in the shipping box):

- Original hard-copy isodose distributions in the axial and sagittal planes through the target volume.  Please ensure that each plane fills an entire page and that a scale is printed on the page.

  - The axial  plane is essentially the central plane and contains the axial film.

  - The film in the sagittal plane is easily seen in the CT scans.

- A completed IROC Houston Head and Neck Phantom-Institution Information form.

- A copy of  results of all film and ion chamber QA measurements.

The following information is required to be submitted to the ITC (results will not be analyzed until the digital treatment plan is received):

- The digital treatment planning data in the RTOG  Data Exchange or DICOM format using either FTP or tape (see the ITC web site for details).

Please cc the IROC Houston (IROCHouston@mdanderson.org) when you inform the ITC of the electronic submission.  This will allow us to keep better track of who has and has not submitted the benchmark electronically.

- Original hard copy isodose distributions in the axial and sagittal plane through the target volume (identical to those sent to the IROC Houston)

- A  copy of the completed IROC Houston Head and Neck Phantom-Institution Information form sent to the IROC Houston

- Send the hard copy data (isodoses and forms) to:

Bill Straube, M.S.

Image Guided Therapy Center

Washington University

4511 Forest Park Ave, Suite 200

St Louis, MO 63108

The Phantom should be imaged, planned, and irradiated as if it were an actual protocol patient, incorporating all of your customary quality assurance checks.

DOSE PRESCRIPTION:

The doses to be delivered to the phantom are a factor of 10 less than the protocol dose specifications, namely:

- Primary PTV.

  - 6.6 Gy to at least 95% of the PTV and

  - < 1% of the PTV receives < 93% of the prescribed dose.

- Secondary PTV (Node or Salivary gland):

  - 5.4 Gy to at least 95% of the PTV and

  - < 1% of the PTV receives < 93% of the prescribed dose.

- Organ at risk:

  - < 4.5 Gy, maximum dose.

- Normal tissue:

  - ( 110% of the prescribed dose (6.6 Gy).

IRRADIATING THE PHANTOM

- Material included in box:

Head Phantom, with 4 TLD capsules taped to the ears.

Phantom insert

Rubber hose

Envelope with background film (hidden from your view; please don’t try to find it)

Pillbox to accept TLD from phantom ears.

Mailing label to return case to IROC Houston at IROC Houston cost.

Traditional IROC Houston TLD block and irradiation table. (Please irradiate this at the time you irradiate the phantom.)

- Adjust the head so that the axial film is perpendicular to the table.

- The adjustment screws in the back of the phantom base should make the phantom adaptable to most head holders.

Procedures:

- Call the IROC Houston with the date that you will irradiate the phantom.  Ask for Nadia Hernandez or leave a message.  Phone number: (713) 745-8989.

- Fill with water:

  - Remove both PVC plugs from the base of the phantom.

  - Thread the rubber hose into one of the filler holes.

  - Fill slowly with water (the rubber hose stretches over most faucets).  You may need to jiggle the head to release air trapped in the nose etc.

  - Remove hose and replace PVC plugs.

- Allow the phantom to sit with water in it for 20 min. to check for leaks.

- Look in the insert space and check for water leakage. If you find any water call the IROC Houston. If not, proceed to the next step.

- This phantom has only one insert.  The same insert is used for both imaging and for treatment. Position the insert.  NOTE: align the two red arrows.  This places the "organ at risk" posterior to the primary target, and the secondary target to the right (phantom #1 and 7 to 25) or left (other phantoms) of the primary target.

- Rotate the insert holders (with the green screws).

- Make sure that there are two TLD capsules taped into each ear.  If they have come out, please tape them back in.  They will remain in for the imaging process, then be removed so as to determine background for the Therapy TLD in the insert.

- CT Scan the phantom as you would a patient. You may wish to scan with 1.5 mm slices especially near the center to better identify the TLD capsules.  Rotate the adjustment screws to support and position the phantom.

- REMOVE THE TLD CAPSULES FROM THE EARS.  Place in a pillbox labeled “ear TLD”.

- Remove insert from head during planning process.  Store insert in a dry place where it will not be irradiated.

- Segment the phantom images contouring the skin, primary and secondary planning target volumes (PTVs) and the organ at risk  (OAR) analog (posterior to the primary target volume crescent) and all 8 TLD volumes.   TLD are in the locations shown in the diagram superior and inferior to the axial film.  Please use the following names for your contours:

PTV_66 for the 1( PTV

PTV_54 for the 2( PTV

CORD for the OAR

66_Sant_TLD and 66_Iant_TLD for the superior and inferior anterior TLDs in the 1( PTV

66_Spost_TLD and 66_Ipost_TLD for the superior and inferior posterior TLDs in the  1( PTV

SCORD_TLD and ICORD_TLD for the superior and inferior TLDs in the OAR

S54_TLD and I54_TLD for the superior and inferior TLDs in the 2( PTV

- The dimensions of the TLD volume are approximately 10 mm long by 2 mm diameter

- The outside dimensions of the TLD capsules are 15 mm long by 4 mm diameter, the TLD axis lies in a sagittal plane. (Both the capsules and the TLD should be visible on CT image)

- Plan the treatment as specified in the DOSE PRESCRIPTION above.

- Perform your customary QA of the IMRT plan prior to irradiating the phantom.

- Irradiate the IROC Houston TLD block according to the instructions provided.

- Look in the insert space and check for water leakage. If you find any water call the IROC Houston. If not, load the insert into the head.

- Treat the phantom with the developed plan as you would an actual patient. Try to avoid positioning the axial film at the abuttment of adjacent MLC leaves or adjacent arcs.  Abutting fields or leaves on the film may increase the uncertainty of the measurement.

- Drain the water.

- Remove the insert and place it in the box.

- Please verify that there is no water in the insert space. If you find any water call the IROC Houston.

- Screw in the two water plugs. Leave the plugs loose.

- Put the empty phantom in the box.

- Make sure that the "ear TLD" pillbox and the rubber hose are in the box.

- Include the dosimetry data discussed above. Complete the attached forms. Be sure to include the scale used on the images coming form your TPS. Isodose lines should include at least the following: 6.6, 5.4, 5, 4.5, 4, 3.5 Gy.

- Return the complete package to the IROC Houston.

###### IROC Houston Head and Neck Phantom- Institution Information

(Original to IROC Houston, copy to ITC)

Please email the IROC Houston , at IROCHouston@mdanderson.org. to let us know when you are going to irradiate the phantom.  We will irradiate TLD standards to meet your schedule.  Please include on the subject: irradiation of IMRT head and neck phantom.

Institution:

Address:

Person performing irradiation:

Physicist to receive report:

Oncologist to receive report:

Person to call in case of questions:

Phone Number:	Fax Number:

Email address: _________________

## Treatment Unit:

Manufacturer:	Model:

In-house specification:

Photon Energy   Nom ________(MV)    IR (TMR 20/TMR 10): ______   %dd(10)x _______

## Intensity Modulation Device:

___ MIMIC                   ___ Multileaf Collimator                  ___ Solid Attenuator Modulation

## IMRT Technique:

___ Segmental (step and shoot) MLC                                  ___ Dynamic MLC

___ Intensity Modulated Arc Therapy (IMAT)

Other: _____________________________________________________________________

Please enclose original copies of your treatment plans.  Include the slices where the films are and scaling factors. FTP the digital treatment plan.

Treatment Planning System:

Manufacturer:_______________________________          Model:_____________________

Software: ____________________________ Version Number:________________________

Treatment of Phantom:

Date of Irradiation:

### Dose specified is to:	___ Muscle 	___ Water

Indicate the dose delivered to the TLD as determined by your treatment planning computer

| TLD | Mean Dose(Gy) | Min Dose(Gy) | Max Dose(Gy) |
| --- | --- | --- | --- |
| Ant 1( PTV superior |  |  |  |
| Ant 1( PTV inferior |  |  |  |
| Post 1( PTV superior |  |  |  |
| Post 1( PTV inferior |  |  |  |
| 2( PTV superior |  |  |  |
| 2( PTV inferior |  |  |  |
| Organ at risk superior |  |  |  |
| Organ at risk inferior |  |  |  |

Results of IMRT QA (please attach copies of measurement data): _________________________

Did you adjust the MU based on these results?___________ If so, how much?_______________

Did you irradiate the phantom in service or clinical mode? ______________________________

Did you use your record and verify system? __________________________________________

Attach copies of the treatment plan including a plan in the axial and sagittal film planes

Comments:

| For Office Use Only | Batch TLD | Phantom ID # | Code | Date Sent | Date Rec'd |
| --- | --- | --- | --- | --- | --- |
|  | Film |  |  |  |  |  |  |

This is a cross sectional view of the insert. The TLD superior and inferior to the axial film.

Note: You need to deliver 6.6 Gy to the 1oPTV (in 1 or more fraction).

Total dose to the 10PTV 6.6 Gy

# Thanks

# Phantom team @ IROC Houston

# Text boxes

Sagittal Film Plane

TLD

TLD

Organ at risk

Secondary PTV

Anterior TLD

Right

Posterior

Left

PosteriorTLD

Primary PTV

# Axial Film Plane

Anterior