## IROC Houston Lung Phantom

# Stereotactic Body Radiation Therapy

Guidelines for Planning and Irradiating the IROC Houston Lung Phantom.

Revised December 2007

Credentialing for this protocol requires four steps: (1) submission of the Facility Questionnaire with supporting documentation to the Advance Technology Consortium (ATC, http://atc.wustl.edu),  (2) a successful dry run test, (3) completion of the phantom treatment experiment and (4) submission of the treatment plan for the first patient treated at the site on this protocol prior to delivering any protocol treatment. The purpose of steps (2) and (3) is to confirm that the dose distribution planned by each institution can be delivered by that institution, and treatment plans can be correctly submitted to the ATC.

The RTOG is requesting that each institution keep the phantom for no more than 1 week.  During this two-week period, the institution will image, plan, and irradiate the phantom and return it to the IROC Houston QA Center. Thank you for your cooperation with this constraint.

This phantom has been designed and constructed by the IROC Houston. The IROC Houston phantom contains an imaging and dosimetric insert. The insert, which is part of the left lung, contains a centrally located GTV (3 cm x 5 cm). There are three orthogonal sheets of radiochromic film passing through the center of the target and two TLD capsules within 0.5 cm of the center of the target. The phantom also contains normal structures: the right lung, the heart, with one TLD capsule in its center and the spinal cord, with one TLD in its center.

If you have any questions, please contact the appropriate person.

IROC Houston	Paola Alvarez	(713) 745-8989    palvarez@mdanderson.org

IROC Houston	Andrea Molineu	(713) 745-8989    amolineu@mdanderson.org

IROC Houston	Geoff Ibbott	(713) 745-8989	gibbott@mdanderson.org

MD Anderson	Mike Gillin	(713) 563-2507	mgillin@mdanderson.org

ITC		Bill Straube	(314) 362-9762	itc@castor.wustl.edu

ITC		Jim Purdy 	(314) 362-2639	purdy@castor.wustl.edu

DOSIMETRY INFORMATION TO BE SUBMITTED:

The following information is to be submitted to the IROC Houston (include in the phantom shipping box):

- Original hard-copy of the plan and isodose distributions without applying correction for tissue heterogeneity in the sagittal, axial and coronal planes through the center of the target volume.  Please ensure that each plane fills an entire page and that a scale is printed on the page.

- Original hard-copy of the plan and isodose distributions applying correction for tissue heterogeneity in the sagittal, axial and coronal planes through the center of the target volume.  Please ensure that each plane fills an entire page and that a scale is printed on the page.  Be sure to apply on each field the number of monitor units obtained from the plan generated without correction for tissue heterogeneity.

- A completed IROC Houston Lung Phantom Institution Information form.

The following information is required to be submitted to the ATC (see protocol for additional submissions) (results will not be analyzed until the digital treatment plan is received):

- The digital treatment planning data, with and without applying correction for tissue heterogeneity, in the RTOG Data Exchange format using either FTP or tape (see the ATC web site for details).

Please cc the IROC Houston (IROCHouston@mdanderson.org) when you inform the ATC of the electronic submission.  This will allow us to keep better track of who has and has not submitted the benchmark electronically.

- Original hard copy isodose distributions without applying correction for tissue heterogeneity in the sagittal, axial and coronal plane through the target center (identical to those sent to the IROC Houston)

- A copy of the completed IROC Houston Lung Phantom-Institution Information form that was sent to the IROC Houston

- Send the hard copy data (isodoses and forms) to:

Bill Straube, M.S.

Image Guided Therapy Center

Washington University

4511 Forest Park Ave, Suite 200

St Louis, MO 63108

DOSE PRESCRIPTION:

DO NOT use correction for tissue heterogeneity when planning and calculating M.U.

Only photon beams with nominal accelerating potential between 4 and 10 MV are allowed.

When static beams are used, a minimum of 7 non-opposing beams should be used.  For arc rotation techniques, a minimum of 340 degrees should be utilized.

Field aperture size and shape should correspond nearly identically to the projection of the PTV along a beam’s eye view.

The prescribed dose to the phantom is 6 Gy to the isodose line circumscribing the PTV. It should be delivered in 1 fraction with the following constraints:

- PTV:

  - CTV = GTV. PTV = GTV + 0.5 cm in axial plane + 1 cm in longitudinal plane.

  - Prescribed dose of 6 Gy to at least 95% of the PTV

  - Minimum dose of 5.4 Gy to at least 99% of the PTV,

  - Hotspots must be manipulated to occur only within the PTV,

  - Maximum dose to any point  ( 2cm away from PTV ( 3.5 Gy

  - Prescription isodose surface must be ( 60% and (  90% of the maximum dose in the plan,

  - Ratio (prescribed isodose vol. / PTV) < 1.2

  - Ratio (50% prescribed isodose vol.) / PTV < 3.6

- Critical Normal Structures (spinal cord, heart, lungs):

  - Constraints over the normal structures are specified in the following table

| Normal structure | Volume | Dose |
| --- | --- | --- |
| Spinal Cord | Any point | ( 1.8 Gy |
| Heart | Any point | ( 3 Gy |
| Whole Lung (Right & Left) | < 10% total vol. | May exceed 2. Gy |

IRRADIATING THE PHANTOM

- Material included in box:

Lung Phantom, with 6 TLD capsules taped to the shell (2 on left side, 2 on right side and 2 on anterior side)

Dosimetric/Imaging insert

Rubber hose

Two acrylic cylinders containing TLD in one of the ends,

Envelope with background film and TLD (hidden from your view; please don’t try to find it)

Mailing label to return case to IROC Houston at the IROC Houston’s expense.

Traditional IROC Houston TLD block and irradiation table. (Please irradiate this at the time you irradiate the phantom.)

Procedures:

- Call the IROC Houston with the date that you expect to irradiate the phantom.  Ask for Nadia Hernandez or leave a message.  Phone number: (713) 745-8989.

- Fill the phantom with water:

  - Thread the rubber hose into the filler hole placed on the base of the phantom.

  - Fill slowly with water (the rubber hose stretches over most faucets).  There is a breathing hole on the phantom, make sure it is open, to allow the pressure to release. You may need to jiggle the phantom to release air trapped inside the cavity.

  - Remove hose and replace acrylic screw.

- Allow the phantom to sit with water in it for 20 min. to check for leaks.

- Look in the insert space and check for water leakage. If you find any water call the IROC Houston. If not, proceed to the next step.

- Position the insert. The end labeled “bottom of insert” should be inserted first. Align the red and green marks.  Make sure that the insert is in its correct position by making small rotations of the insert around its central axis. When it is in the correct position it will be locked in place by an indentation at the end of the insert.

- Position the acrylic cylinder labeled “spinal cord cylinder” in the hole labeled “spinal cord”. The hole and the cylinder are marked in blue. You will see a TLD capsule in the cavity closed with a screw.  The end with the TLD should be inserted first.

- Position the acrylic cylinder labeled “heart cylinder” in the hole labeled “heart”. The hole and the cylinder are marked in green. You will see a TLD capsule in the cavity closed with a screw.  The end with the TLD should be inserted first.

- Position and CT the phantom as you would a patient including immobilization techniques. You may wish to scan with 1.5 mm slices especially near the target to better identify the TLD capsules. NOTE: There are TLD on the external shell of the phantom to give us an estimate of the CT dose to the target.

- Remove insert from the phantom during planning process. Store insert in a dry place. Store insert and phantom where they will not be irradiate.

- Segment the phantom images contouring the skin, lungs, heart, spinal cord and PTV. Note that the CTV = GTV. PTV = GTV + 0.5 cm in axial plane + 1 cm in longitudinal plane. Also contour all the 4 TLD volumes.  Please use the following names for these contours:

PTV_TLD_sup for the superior TLD in the

target,

PTV_TLD_inf for the inferior TLD in the

target,

HEART_TLD for TLD in the heart

CORD_TLD for the TLD in the spinal cord

( The dimensions of the TLD volume are approximately 10 mm long by 2 mm diameter

( The outside dimensions of the TLD capsules are 15 mm long by 4 mm diameter; the TLD axis is normal to the axial plane. (The capsules and the TLD should be visible on CT image)

- Plan the treatment as specified in the DOSE PRESCRIPTION above.

- Look in the insert space and check for water leakage. If you find any water call the IROC Houston. If not, follow the instructions in step 5 to position the insert again and proceed to the next step.

- Perform your customary QA of the plan prior to irradiating the phantom.

- Irradiate the IROC Houston TLD block according to the instructions provided.

- Position the phantom as you would a protocol patient including immobilization techniques. Try to avoid positioning the axial film at the abutment of adjacent MLC leaves or adjacent arcs.  Abutting fields or leaves on the film may increase the uncertainty of the measurement.

- REMOVE THE TLD CAPSULES LOCATED ON THE EXTERNAL SHELL. Put them into the designated envelope.

- Irradiate the phantom with the developed plan.

- Remove the insert and place it in the box.

- Remove the acrylic cylinders from holes and place them in the box.

- Please verify that there is no water in the insert space. If you find any water call the IROC Houston.

- Remove the screw on the base of the phantom and drain the water from the phantom.

- Put the empty phantom in the box.

- Make sure that the envelope with the TLD on the shell and the rubber hose are in the box.

- Include the dosimetry data discussed above. Complete the attached forms. Be sure to include the scale used on the images coming from your TPS.

- Return the complete package to the IROC Houston.

###### IROC Houston Lung Phantom Institution Information

(Original to IROC Houston, copy to ITC)

Please email the IROC Houston , at IROCHouston@mdanderson.org. to let us know when you are going to irradiate the phantom.  We will irradiate TLD standards to meet your schedule.  Please include on the subject: irradiation of  lung phantom.

Institution:

Address:

Person performing irradiation:

Person to receive report:

Oncologist to receive report:

Person to call in case of questions:

Phone Number:	Fax Number:

Email address: _________________________________________________________________

## Treatment Unit:

Manufacturer:	Model:

In-house specification:

Photon Beam:

Nom. accelerating potential:________(MV)   IR (TMR 20/TMR 10): _____  %dd(10)x ___________

Collimator :

MLC (       MIMIC (      Other:___________________ No. of leaves: _________________

Stereotactic System (if modification to linac):______________________________________

Manufacter:____________________________Model:____________________________

Other: _____________________________________________________________________

Please enclose original copies of your treatment plans.  Include the coronal, axial and sagittal planes through the target center. Include scaling factors for each plane. FTP the digital treatment planning data.

Treatment Planning System:

Manufacturer: _______________________________          Model:_____________________

Software: ____________________________ Version Number:________________________

Treatment of Phantom:

Date of Irradiation:

### Dose specified is to:	( Muscle 	( Water

Indicate the dose delivered to the TLD as determined by your treatment planning computer

| TLD | Mean Dose (Gy) | Min. Dose (Gy) | Max. Dose (Gy) |
| --- | --- | --- | --- |
|  | Hetero correction on | Homo correction off | Hetero correction on | Homo correction off | Hetero correction on | Homo correction off |
| PTV_TLD_sup |  |  |  |  |  |  |
| PTV_TLD_inf |  |  |  |  |  |  |
| HEART_TLD |  |  |  |  |  |  |
| CORD_TLD |  |  |  |  |  |  |

Results of the QA:______________________________________________________________

_____________________________________________________________________________

Did you adjust the MU based on these results?_______________If so, how much?___________

Attach copies of the treatment plan including slices in the sagittal, axial and coronal film planes.

Comments:

| For Office Use Only | TLD Batch | Film Batch | Phantom ID # | Code | Date Sent | Date Rec'd |
| --- | --- | --- | --- | --- | --- | --- |

This is a cross sectional view of the phantom.

Note: Please ignore all markings on the external shell of the phantom, use your own system to position the phantom.

Note: You need to deliver 6.0 Gy to the PTV (in 1 or more fraction).

Total dose to the PTV 6.0 Gy

# Thanks

# Phantom team @ IROC Houston

# Text boxes

( Position of TLD

Axial Film Plane

The phantom should be imaged, planned and irradiated as if it were an actual protocol patient, incorporating all of your customary quality assurance checks.

Sagittal Film Plane

Coronal Film Plane

GTV

Heart

Spinal

Cord

Left

Lung

Right

Lung

Right

Left

Posterior

Anterior