## IROC Houston Prostate Phantom

Guidelines for Planning and Treating the IROC Houston IMRT Prostate Phantom.

Revised December 2007

Credentialing for this protocol requires three steps: (1) submission of the Facility Questionaire with supporting documentation to the Image Guided Therapy Center (ITC, http://itc.wustl.edu), (2) a successful dry run test, and (3) completion of the phantom treatment experiment. The purpose of steps (2) and (3) is to confirm that the IMRT dose distribution planned by each institution can be delivered by that institution, and correctly submitted to the ITC.

The RTOG is requesting that each institution keep the phantom for a period of time no more than 2 weeks.  During this two-week period, the institution will image, plan, and treat the phantom and return it to the IROC Houston QA Center. Thank you for your cooperation.

This phantom has been designed and constructed by the IROC Houston. The IROC Houston phantom contains two inserts. The water fillable imaging insert contains the CTV (prostate) and critical normal structures (bladder and rectum). The dosimetric insert contains TLD at 2 locations and perpendicular sheets of film in order to evaluate the dose to the center of the prostate. There are 2 more TLD to evaluate the dose to each femoral head.

If you have any questions, please contact the appropriate person.

IROC Houston	Geoff Ibbott	(713) 745-8989	gibbott@mdanderson.org

IROC Houston	Paola Alvarez	(713) 745-8989    palvarez@mdanderson.org

IROC Houston	Andrea Molineu	(713) 745-8989    amolineu@mdanderson.org

MD Anderson	Mike Gillin	(713) 563-2507	mgillin@mdanderson.org

ITC		Jim Purdy 	(314) 362-2639	purdy@castor.wustl.edu

ITC		Bill Straube	(314) 362-9762	itc@castor.wustl.edu

DOSIMETRY INFORMATION TO BE SUBMITTED:

The following information is to be submitted to the IROC Houston (include it in the shipping box):

- A completed IROC Houston Prostate Institution Information form.

- Original hard copy of the plan and isodose distributions in the sagittal and coronal plane through the target center.

The following information is required to be submitted to the ITC (results will not be analyzed until the digital treatment plan is received):

- The digital treatment planning data in the RTOG Data Exchange format using either FTP or tape (see the ITC web site for details).

Please cc the IROC Houston (IROCHouston@mdanderson.org) when you inform the ITC of the electronic submission.  This will allow us to keep better track of who has and has not submitted the benchmark electronically.

- Original hard copy isodose distributions in the sagittal and coronal plane through the target center (identical to those sent to the IROC Houston)

- A copy of the completed IROC Houston Prostate Phantom-Institution Information form that was sent to the IROC Houston

- Send the hard copy data (isodoses and forms) to:

Bill Straube, M.S.

Image Guided Therapy Center

Washington University

4511 Forest Park Ave, Suite 200

St Louis, MO 63108

DOSE PRESCRIPTION:

The total dose to the phantom is 6 Gy, with the following constraints:

- Prostate PTV.

  - Total dose of 6 Gy to at least 98% of the PTV and

  - A maximum dose of 6.4 Gy may be given to < 2% of the PTV. This maximum dose volume must not be shared by critical normal structures.

- Critical Normal Structures (bladder, rectum and femoral head):

  - No part of these normal organs shall receive more than 6.7 Gy

  - Constraints over the normal structure are specified in the following table

IRRADIATING THE PHANTOM

- Material included in box:

Prostate Phantom

Dosimetric insert

Imaging insert

Two plastic screws

Rubber hose

Two acrylic cylinders  labeled  “imaging cylinders”

Two acrylic cylinders containing TLD in one of the ends, labeled “dosimetric cylinders”

Envelope with background film (hidden from your view; please don’t try to find it)

Mailing label to return case to IROC Houston at the IROC Houston’s cost.

Traditional IROC Houston TLD block and irradiation table. (Please irradiate this at the time you irradiate the phantom.)

Procedures:

- Call the IROC Houston with the date that you will irradiate the phantom.  Ask for Nadia Hernandez or leave a message.  Phone number: (713) 745-8989.

- Fill the phantom with water:

  - Thread the rubber hose into one of the filler holes placed on the base of the phantom.

  - Fill slowly with water (the rubber hose stretches over most faucets).  You may need to jiggle the phantom to release air trapped inside the cavity.

  - Remove hose and replace acrylic screws.

- Allow the phantom to sit with water in it for 20 min. to check for leaks.

- Fill the imaging insert with water:

  - Remove both PVC plugs from the top of the insert.

  - Thread the rubber hose into one of the filler holes.

  - Fill slowly with water (the rubber hose stretches over most faucets).  You may need to jiggle the head to release air trapped between the different structures.

  - Remove hose and replace PVC plugs.

  - There is a third plug labeled in blue. Remove this plug and use this hole to add more water, if needed, to avoid air bubble inside the insert.

5.	Look in the insert space and check for water leakage. If you find any water call the IROC Houston. If not, proceed to the next step.

6.	Position the imaging insert. The side without plugs and holes should be inserted. Align the two red lines.  This places the rectum posterior to the prostate. Make sure that the insert is in its correct position by making small rotations of the insert around its central axis. When it is in the correct position it will be locked in place by an indentation at the end of the insert.

7.	Position one of the “imaging cylinders” in each one of the holes in the femoral head.

- Position and CT the phantom as you would a patient. You may wish to scan with 1.5 mm slices especially near the center to better identify the edges of each organ. Remove the imaging insert. Drain the water from the insert. Place the insert in the box.

- Segment the phantom images contouring the skin, prostate, bladder and rectum.

- Look in the insert space and check for water leakage. If you find any water call the IROC Houston. If not, proceed to the next step.

- Position the dosimetric insert. The side without a handle should be inserted.  Align the two red lines.  As in step 6 be sure that the insert is in its correct position.

- Position the “dosimetric cylinders” in each one of the holes in the femoral head following the color code. Insert the side with the screw. You will see a TLD capsule in each one of the cavities closed with the screw.

- Perform your customary QA of the IMRT plan prior to irradiating the phantom. Include in the form values and all the information you consider are relevant for it analysis.

- Irradiate the IROC Houston TLD block according to the instructions provided.

- Treat the phantom with the developed plan as you would a protocol patient.

- Remove the dosimetric insert and place it in the box.

- Remove the acrylic legs from holes in the femoral head and place them in the box.

- Please verify that there is no water in the insert space. If you find any water call the IROC Houston.

- Remove the screw on the base of the phantom and drain the water in the phantom.

- Put the empty phantom in the box.

- Make sure that the rubber hose and the plastic screws are in the box.

- Include the dosimetry data discussed above. Complete the attached forms. Be sure to include the scale used on the images coming from your TPS.

- Return the complete package to the IROC Houston.

###### IROC Houston Prostate Phantom Institution Information

(Original to IROC Houston, copy to ITC)

Please email the IROC Houston, at IROCHouston@mdanderson.org. to let us know when you are going to irradiate the phantom.  We will irradiate TLD standards to meet your schedule.  Please include on the subject: irradiation of IMRT prostate phantom.

Institution:

Address:

Person performing irradiation:

Physicist to receive report:

Oncologist to receive report:

Person to call in case of questions:

Phone Number:	Fax Number:

Email address: _________________________________________________________________

## Treatment Unit:

Manufacturer:	Model:

In-house specification:

Photon Energy   Nom ________(MV) IR (TMR 20/TMR 10): ______   %dd(10)x _______

## Intensity Modulation Device:

( MIMIC                   ( Multileaf Collimator                  ( Solid Attenuator Modulation

## IMRT Technique:

( Segmental (step and shoot) MLC     ( Dynamic MLC   ( Intensity Modulated Arc Therapy (IMAT)

Other: _____________________________________________________________________

Please enclose original copies of your treatment plans.  Include the coronal and saggital planes through the target center. Include scaling factors for each plane. FTP the digital treatment plan.

Treatment Planning System:

Manufacturer: _______________________________          Model:_____________________

Software: ____________________________ Version Number:________________________

Treatment of Phantom:

Date of Irradiation:

### Dose specified is to:	( Muscle 	( Water

Indicate the dose delivered to these specific points as determined by your treatment planning computer

| Point | Dose (Gy) |
| --- | --- |
| Center of the prostate |  |
| TLD position on femoral head (Left)* |  |
| TLD position on femoral head (Right)* |  |

* Dose to the center of the TLD position on the femoral head, on the axial plane through the center of the prostate.

Results of IMRT QA:____________________________________________________________

______________________________________________________________________________

Did you change the M.U. based on your QA? (No  (Yes

Attach copies of the treatment plan including slices in the sagittal and coronal film planes.

Comments:

| For Office Use Only | TLD Batch | Film Batch | Phantom ID # | Code | Date Sent | Date Rec'd |
| --- | --- | --- | --- | --- | --- | --- |

This is a cross sectional view of the phantom with the imaging insert in place.

Notes:

- You need to deliver 6 Gy to the PTV (in 1 or more fraction).

- Please ignore all markings on the external shell of the phantom, use your own system to position the phantom.

# Text boxes

Sagittal Film Plane

femoral

head

Position for

TLD

| Normal Organs | No more than 15% volume receives dose that exceeds | No more than 25% volume receives dose that exceeds | No more than 35% volume receives dose that exceeds | No more than 50% volume receives dose that exceeds |
| --- | --- | --- | --- | --- |
| Bladder | 6.7 Gy | 6.3 Gy | 6.0 Gy | 5.7 Gy |
| Rectum | 6.3 Gy | 6.0 Gy | 5.7 Gy | 5.0 Gy |

The phantom should be imaged, planned and irradiated as if it were an actual protocol patient, incorporating all of your customary quality assurance checks.

Position for

TLD

rectum

Right

Posterior

Left

femoral

head

bladder

prostate

Coronal Film Plane

Anterior