The University of Texas M. D. Anderson Cancer Center1515 Holcombe Boulevard, Unit 547Houston, Texas 77030

REPORT TO THE AAPM THERAPY PHYSICS COMMITTEE

Report No. 125 August 1, 2006 -October 30, 2006

a) Administrative

TG-51 CONVERSIONS

There have been no changes to senior staff at the RPC

1400

during this reporting period. A TLD technician,Stephanie Lampe, has moved from the RPC to the

1200

ADCL, to replace another ADCL staff member who left

1000

recently. Ms. Lampe had worked previously at the

800

ADCL, but returns with increased responsibilities. This move, in the wake of Dr. Hanson’s retirement in

600

September, positions the ADCL to respond efficiently

400

to requests from customers, but maintains the expenses of operation at a reasonable level.

200

0

b) Studies and Results

The RPC presently monitors 1,474 megavoltage therapy sites in North America, Europe, and elsewhere

in the world, that participate in cooperative group • In 2003, the RPC began sending TLD to institutions on

an annual schedule, rather than waiting six months from

receipt of the previous shipment. This has been a

successful change in that it allows us to manage our

TLD mailings better throughout the year. In addition, clinical trials funded by the NCI. The cooperative

groups monitored include ACOSOG, ACRIN, CALGB,

COG, ECOG, GOG, NABTT, NCCTG, NSABP,

RTOG and SWOG. We also communicate regularly

with the Cancer Trials Support Unit (CTSU) to assure

that institutions participating through their programs are this change, along with more aggressive pursuit of

delinquent institutions, has reduced the TLD return time

from about 100 days at the beginning of 2001, to about properly monitored. 60 days today.

• On-Site Dosimetry Reviews: In the year since October

Chart Title

1, 2005, RPC physicists have performed on-site

140

dosimetry evaluation at 21 institutions, evaluating 190beams (a “beam” is a single photon beam or a cadre of

120

electron beams). Twenty (95%) of the institutions

100

received at least one recommendation for actions that

DAYS#)Poly. (#DAYS

# DAYS

80

should improve their quality assurance programs.

60

During the last three years, additional procedures

40

including measurement of small fields were added, to

help evaluate IMRT.

20

TLD: Between October 1, 2003 and September 30

0

•

2006, 28,954 (distinct energies) beams were monitored

DATE

with TLD at the monitored institutions. Overall, 1.8% of the irradiated TLD received doses that disagreed with the institution’s stated dose by more than 5%.

• Patient Treatment Review: In the year since October Institutions currently using the obsolete calibration

1, 2005, individual protocol patient treatment records protocol (TG-21) exhibited a higher discrepancy rate

were evaluated for 246 patients treated on GOG, than the institutions that had converted to the new

NSABP, NCCTG, and RTOG protocols. Of these, 44 calibration protocol (TG-51). The number of

patients received brachytherapy treatments. institutions still using TG-21 has decreased steadily

• Credentialing Processes: The RPC participates in the over the past few years. As of October 1, 2006, 85% of credentialing of institutions for protocols involving the institutions we monitor have converted to TG-51.

American Association of Physicists in MedicineThe Association’s Scientific Journal is Medical PhysicsRadiological Physics Center: Telephone (713) 745-8989 Fax (713) 794-1364 https://irochouston.mdanderson.org

advanced technologies including brachytherapy, IMRT, stereotactic radiosurgery (SRS) and stereotactic body radiation therapy (SBRT). This activity is partially supported by a subcontract from the Advanced Technologies for Clinical Trials grant. For IMRT and SRS, credentialing includes irradiation of an anthropomorphic phantom provided by the RPC that contains anatomic structures and dosimeters. To date, the RPC has issued 244 reports to institutions that irradiated the RPC head and neck IMRT phantom. Of these, 173 indicated that the irradiation complied with the institution’s treatment plan within criteria agreed to by the RPC and the RTOG. The first-time pass rate was 69%. The RPC’s lung phantom has been used to credential institutions to participate in a stereotactic lung trial which has recently been completed. Twenty-two reports were issued. Of them, 15 indicated that the irradiation complied with the institution's treatment plan within criteria agreed to by the RPC and RTOG. Valuable data regarding heterogeneity corrections has been obtained and is being prepared for publication. The RPC plays the lead role in credentialing institutions for a partial breast irradiation (PBI) trial run jointly by NSABP and RTOG. Credentialing requires completion of facility and knowledge-assessment questionnaires (using web-based forms) and electronic submission of a benchmark treatment plan. To date, 1757 applications for credentialing have been received (all or part) and 1394 credentials have been issued (3D CRT arm-792; Mammosite arm-497; multicatheter arm-115). The remainder are awaiting submission of required information or replies to questions. Complete applications are processed within 2 business days.

Low-Energy Brachytherapy Sources: Dr. Ibbott represents the RPC on a subcommittee appointed by the AAPM to address the use of new sources. The RPC acts as a clearinghouse of information and makes available on a web site a list of sources meeting the AAPM dosimetric prerequisites. Preliminary discussions have been held with ESTRO physics committee regarding collaboration on the RPC website.

Planning Workstation: An educational grant from Varian Corporation enabled the RPC to acquire an Eclipse treatment planning workstation. RPC employees have received training, and the workstation has been commissioned. The TPS will be used to re-calculate institutions' treatment plans for verification.

Database: The RPC database continues to greatly facilitate our work. Results of remote measurements, on-site dosimetry reviews, and treatment record reviews are available to RPC staff immediately.

Webpage: The RPC webpage has undergone a major revision. We have now implemented online web forms

for a number of credentialing activities, including RTOG 0413/NSABP B-39.

AAPM Oversight: The AAPM Therapy Physics Committee continues to be our scientific advisory committee. We report to this committee three times per year, and participate in subcommittees, task group, and working groups. A task group of the TPC performs a one-day in depth evaluation annually. The next meeting of the TPC will be held in Houston in March, other than the 12th to 16th .

Clinical Advisory Committee: A group of 5 radiation oncologists was formed as a clinical advisory committee. The Committee is contacted when questions arise regarding RPC operations.

c)	Plans

Plans for this budget period are not significantly changed from those outlined in the application for the current grant cycle.

On-Site Dosimetry Reviews: Each FTE traveling physicist should review 50 beams.

Mailed TLD Program: Our criteria for acceptability have not been changed. The use of the existing TLD system for dedicated SRS equipment is being evaluated.

Credentialing: We will continue to work with the groups to focus our efforts appropriately.

Patient Treatment Review: Except for some studies evaluated by the RTOG QA office, the RPC is the only QA office that focuses on technical evaluation of radiation dose (dosimetry review). The RPC will continue to review some fraction of patients to assess the quality of the data currently being submitted. The Eclipse workstation will be incorporated into our review of patient records, to replace in-house developed software used for recalculation of dose distributions.

Anthropomorphic Phantoms: Additional phantoms have been constructed and additional phantoms have been requested, to meet the demand for credentialing on all existing protocols and newly developed protocols. Special attention will be given to the evaluation of heterogeneity corrections in lung protocols. Changes in priorities at the MD Anderson machine shop have required that we eliminate our support of a machinist position, and use the funds to pay for construction of phantoms at an outside facility.

Liaison with Cooperative Groups, AAPM, etc.: The AAPM Therapy Physics Committee continues as our scientific advisory body. The RPC participated in the development and conduct of a roundtable conference on quality assurance practices in radiation oncology held at

the NCI in September 2005. This workshop was in preparation for a larger workshop now scheduled to be held in the Spring of 2007. The RPC is playing a lead role in organizing and conducting this workshop.

Data Transfer: The RPC, through the ATC subcontract, continues to develop and implement electronic data exchange capabilities. We now use electronic data routinely in the evaluation of phantom irradiations.

Webpage: We continue to add capabilities to the RPC webpage, and plan in the near future to convert our TLD review mechanism to a fully electronic process. Discussions have been held with the AAPM regarding access to the RPC database for data mining activities.

PARTICIPANT FEE: Institutions invoiced FY07 1426 No XRT/Canceled/Inactive 7 Invoiced by RDS 2 Institutions paid 778

PUBLICATIONS AND ABSTRACTS Publications Accepted/Published (2001-2006):

Kirsner, S.M., Prado, K.L., Tailor, R.C., and Bencomo, J.A.: Verification of the accuracy of 3D calculations of breast dose during tangential irradiation: measurements in a breast phantom. J. Applied Clin. Med. Phys., 2 (3), pp. 149-156, 2001.

Melia, Sc.M., Michele; Abramson, M.D., David; Albert, M.D., Daniel; Boldt, M.D. Culver; Earle, M.D., John; Hanson, Ph.D., William; Montague, Paul; Moy, Ph.D., Claudia; Schachat, M.D., Andrew; Simpson, M.D., Rand; Straatsma, M.D., Bradley; Vine, M.D., Andrew; and Weingeist, M.D., Ph.D., Thomas: Collaborative Ocular Melanoma Study (COMS) Randomized Trial of I-125 Brachytherapy for Medium Choroidal Melanoma I. Visual Acuity after 3 Years, COMS Report No. 16”, Opthalmology, 108 (2):348-366, 2001.

Kim, C-H., Reece, W. D., and Cho, S. H. Computer simulation of radiation exposure in a S/G channel head. Trans. Am. Nucl. Soc., 84:325-326, 2001.

Tailor, R., Hanson, W., Calculated absorbed-dose ratios, TG-51/TG-21, for most widely used cylindrical and parallel-plate ion chambers over a range of photon and electron energies, Med. Phys. Vol. 29:1464-1472, 2002.

Gifford, K.A., Followill, D.S., Liu, H.H., and Starkschall, G. Verification of the accuracy of a photon dose-calculation algorithm. J. Applied Clin. Med. Phys. 3:26-45, 2002.

Villarruel, S., Ibbott, G.S., and Lai-Fook, S.J.: Effect of concentration and hydration on restriction of albumin by lung interstitium. Microvascular Research 63, 27-40, 2002.

Cadman, P., Bassalow, R., Sidhu, N.P.S., Ibbott, G., Nelson, A. Dosimetric considerations for validation of a sequential IMRT process with a commercial treatment planning system. Physics in Medicine and Biology Vol. 47, 3001-3010, 2002.

Ibbott, G., Beach, M., Maryanski, M. An anthropomorphic head phantom with a BANG® polymer gel insert for dosimetric evaluation of IMRT treatment delivery. Standards and Codes of Practice in Medical Radiation Dosimetry, Proceedings of an International Symposium, Vienna, Vol. 2, pp. 361-368, November 25-28, 2002.

9.	Ibbott, G., Nelson, A., Followill, D., Balter, P., Hanson,

W. An anthropomorphic head and neck phantom for evaluation of intensity modulated radiation therapy. Standards and Codes of Practice in Medical Radiation Dosimetry, Proceedings of an International Symposium, Vienna, Vol. 2, pp. 209-217, November 25-28, 2002.

Izewska, J., Svensson, H., Ibbott, G. Worldwide QA networks for radiotherapy dosimetry. Standards and Codes of Practice in Medical Radiation Dosimetry, Proceedings of an International Symposium, Vienna, Vol. 2, pp. 139-155, November 25-28, 2002.

Aguirre J, Tailor R, Ibbott G, Stovall M, Hanson W. TLD as a remote verification of output for radiotherapy beams: 25 years of experience. Standards and Codes of Practice in Medical Radiation Dosimetry, Proceedings of an International Symposium, Vienna, Vol. 2, pp. 191-199, November 25-28, 2002.

12.	Tailor R, Hanson W, and Ibbott G, TG-51 Experience from 150 institutions, common errors, and helpful hints,

J. Applied Clin. Med. Phys., Vol. 4, pp.102-111, 2003.

Urie, M., FitzGerald, T.J., Followill, D., Laurie, F., Marcus, R., Michalski, J. Current calibration, treatment, and treatment planning techniques among institutions participating in the Children’s Oncology Group. Int. J. of Radiat. Oncol., Biol., Phys. 1:245-260, 2003.

14.	Followill D.S., Hanson, W.F., Ibbott, G.S., Eglezopoulos, L.R., and Chui, C.S. Differences in electron beam dosimetry using two commercial ionization chambers and the TG-21 protocol: another reason to switch to TG-51. J. Applied Clin. Med. Phys.

4: 124-131, 2003.

Nag, S., Quivey, J.M., Earle, J.D., Followill, D.S., Fontanesi, J., and Finger, P. The American Brachytherapy Society Recommendations for Brachytherapy of Uveal Melanomas, Int. J. of Radiat. Oncol., Biol., Phys. 56:544-555, 2003.

Krintz, A.L., Hanson, W.F., Ibbott, G.I. and Followill, D.S., A Reanalysis of the Collaborative Ocular Melanoma Study Medium Tumor Trial Eye Plaque Dosimetry, Int. J. of Radiat. Oncol., Biol., Phys. 56:889-898, 2003.

Followill, D.S., Stovall, M.S., Kry, S.F., and Ibbott, G.S., Neutron source strength measurements for Varian, Siemens, Elekta, and General Electric linear accelerators. J. Applied Clin. Med. Phys. 4:189-194, 2003.

Diener-West, M., Albert, D. M., Frazier Byrne, SI, Davidorf, F. H. Followill, D. S., Green, R.L., Hawkins, B.S., Kaiser, P.K., Robertson, D.M., and Straatsma, B.R., Comparison of Clinical, Echographic and Histopathologic Measurements from Eyes with Medium-Sized Choroidal Melanoma: in the Collaborative Ocular Melanoma Study. COMS Report No. 21, The Collaborative Ocular Melanoma Study Group, Archives of Ophthalmology 121: 2003.

Tailor, R.C., Followill, D.S., Hernandez, N., Ibbott, G.S., and Hanson, W.F., “Predictability of electron cone ratios with respect to linac make and model,” J. Applied Clin. Med. Phys., Vol. 4 (2), pp. 172-178, 2003.

Cho, S.H. and Ibbott, G.S., “Reference photon dosimetry data: A preliminary study of in-air off-axis factor, percentage depth dose, and output factor of the Siemens Primus linear accelerator”, J. Applied Clin. Med. Phys. 4 (4): 300-306, 2003.

Cho S, Reece W, Kim C. Validity of two simple rescaling methods for electron/beta dose point kernels in heterogeneous source-target geometry. Radiation Physics and Chemistry 69:265-72, 2004.

Rivard M, Coursey B, DeWerd L, Hanson W, Huq M, Ibbott G, Mitch M, Nath R, Williamson J. Update of AAPM Task Group No. 43 Report: A revised AAPM protocol for brachytherapy dose calculations. Med. Phys. 31:633-74, 2004.

DeWerd L, Huq M, Das I, Ibbott G, Hanson W, Slowey T, Williamson J, Coursey B. Procedures for establishing and maintaining consistent air-kerma strength standards for low-energy, photon-emitting brachytherapy sources: Recommendations of the Calibration Laboratory Accreditation Subcommittee of the American Association of Physicists in Medicine. Med. Phys. 31:675-81, 2004.

BenComo J, Chu C, Tello V, Cho S, Ibbott G. Anthropomorphic breast phantoms for quality assurance and dose verification. J. Applied Clin. Med. Phys. 5:36-49, 2004.

Cho S, Vassiliev O, Lee S, Liu H, Ibbott G, Mohan R. Reference photon dosimetry data and reference phase space data for the 6 MV photon beam from Varian Clinac 2100 series linear accelerators. Under revision for Med. Phys., 2004.

Followill D, Davis D, Ibbott G. Comparison of Electron Beam Characteristics from Multiple Accelerators. Int. J. of Radiat. Oncol., Biol., Phys. 59:905-10, 2004.

Marcus D, Peskin E, Maguire M, Weissgold D, Alexander J, Fine S, Followill D, the AMDRT Research Group. The Age-Related Macular Degeneration Radiotherapy Trial (AMDRT): One Year Results from a Pilot Study. American Journal of Ophthalmology 138:818-28, 2004.

28. Nag S, Cardenes H, Chang S, Das I, Erickson B, Ibbott G, Lowenstein J, Roll J, Thomadsen B, Varia M. Proposed Guidelines For Image-Based Intracavitary Brachytherapy For Cervical Carcinoma: A Report From The Image-Guided Brachytherapy Working Group. Int.

J. Radiat. Oncol., Biol., Phys. 60:1160-72, 2004.

Olch A, Kline R, Ibbott G, Anderson J, Deye J, FitzGerald T, Followill D, Gillin M, Huq M, Palter J, Purdy J, Urie M. Quality Assurance for Clinical Trials: A Primer for Physicists. AAPM Report No. 86, 2004.

Ibbott G. Applications of Gel Dosimetry. Journal of Physics: Conference Series 3:58-77, 2004.

Gifford K, Horton J, Steger T, Heard M, Jackson E, Ibbott G. Verification of Monte Carlo calculations around a Fletcher Suit Delclos Ovoid with Normoxic

Polymer Gel Dosimetry. Journal of Physics: Conference Series 3:217-20, 2004.

Heard M, Ibbott G. Measurement of brachytherapy sources using MAGIC gel. Journal of Physics: Conference Series 3:221-3, 2004.

Beddar A, Tailor R. Calibration of low energy electron beams from a mobile linear accelerator with plane-parallel chambers using both TG-51 and TG-21 protocols. Phys .Med. Biol. 49: N105-N110, 2004.

Cho S, Vassiliev O, Lee S, Liu H, Ibbott G, Mohan R. Reference photon dosimetry data and reference phase space data for the 6 MV photon beam from Varian Clinac 2100 series linear accelerators. Med. Phys. 32:137-48, 2005.

Molineu A, Followill DS, Balter PA, Hanson WF, Gillin MT, Huq MS, Eisbruch A, Ibbott GS. Design and Implementation of an Anthropomorphic Quality Assurance Phantom for Intensity Modulated Radiation Therapy for the Radiation Oncology Group. Int. J. of Radiat. Oncol. Biol. Phys. 63:577-83, 2005.

Williamson J, Butler W, DeWerd L., Huq M, Ibbott G, Li, Z, Mitch M, Nath R, Rivard M, Todor D. Recommendations of the American Association of Physicists in Medicine regarding the Impact of Implementing the 2004 Task Group 43 report on Dose Specification for 103Pd and 125I Interstitial Brachytherapy. Med. Phys. 32:1424-39, 2005.

Zhang G, Guerrero T, Segars W, Huang T, Bilton S, Lin KP, Ibbott G, Dong L, Forster K. Elastic Image Mapping for 4D Dose Estimation in Thoracic Radiotherapy. In press, Radiation Protection Dosimetry, 2005.

Gifford KA, Horton Jr. JL, Jackson EF, Steger III TR, Heard MP, Mourtada F, Lawyer AA, Ibbott GS. Comparison of Monte Carlo calculations around a Fletcher Suit Delclos ovoid with radiochromic film and normoxic polymer gel dosimetry. Medical Physics 32:2288-94, 2005.

Cho SH. Estimation of tumor dose enhancement due to gold nanoparticles during typical radiation treatments: A preliminary Monte Carlo study. Physics in Medicine and Biology 50:163-73, 2005.

Kry SF, Salehpour M, Followill DS, Stovall M, Kuban DA, White RA, Rosen II. Out-of-Field Photon and Neutron Dose Equivalents from Step-and-Shoot

Intensity-Modulated Radiation Therapy. Int J Radiat Oncol Biol Phys 62:1204-16, 2005.

Kry SF, Salehpour M, Followill DS, Stovall M, Kuban DA, White RA, Rosen II. The Calculated Risk of Fatal Secondary Malignancies from Intensity-Modulated Radiation Therapy. Int J Radiat Oncol Biol Phys 62:1195-1203, 2005.

Halvorsen H, Das IJ, Fraser M, Freedman DJ, Rice III RE, Ibbott GS, Parsai EI, Robin Jr. TT, Thomadsen BR. AAPM Task Group 103 Report on Peer Review in Clinical Radiation Oncology Physics. Journal of Applied Clinical Medical Physics 6:50-64, 2005.

Tailor RC, Hanson WF, Wells N, Ibbott GS. Consistency of absorbed dose to water measurements using 21 ion-chamber models following the AAPM TG51 and TG21 calibration protocols. Med Phys 33:1818-28, 2006.

Randall M, Ibbott GS. Intensity-modulated radiation therapy for gynecologic cancers: pitfalls, hazards, and cautions to be considered. Seminars in Radiat Oncol 16:138-43, 2006.

Briere TM, Tailor R, Tolani N, Prado K, Lane R, Woo S, Ha C, Gillin MT, Beddar S. Patient dosimetry for total body irradiation using single-use MOSFET detectors. Submitted to Medical Physics, 2006.

Minniti R, Chen-Mayer H, Seltzer SM, Huq MS, Bryson L, Slowey T, Micka JA, DeWerd LA, Wells N, Hanson WF, Ibbott GS. The US radiation dosimetry standards for 60 Co therapy level beams, and the transfer to the AAPM accredited dosimetry calibration laboratories. Med Phys 33: 1074-7, 2006.

Ibbott G, Molineu A, Followill D. Independent evaluations of IMRT through the use of an anthropomorphic phantom. Technology in Cancer Research and Treatment 5:481-8, 2006.

Timmerman R, Galvin J, Michalski J, Straube W, Ibbott G, Martin E, Abdulrahman R, Swann S, Fowler J, Choy

H. Multicenter Trials with SBRT in Lung Cancer. Acta Oncologica, 2006 (in press).

Abstracts

1.	Balter P, Lowenstein J, and Hanson W: Electron Calibrations: Parallel Plate Chambers vs. Cylindrical Chambers Using TG-51. Medical Physics, Vol. 28: 1214, 2001.

Radford, D.A., Followill, D.S., and Hanson, W.F.: A Standard Method of Quality Assurance for Intensity Modulated Radiation Therapy of the Prostate. Medical Physics 28 (6): 1211, 2001.

Followill, D.S., Hernandez, N., and Hanson, W.F.: Standard Wedge and Tray Transmission Values for Varian, Seimens, Elekta/Philips Accelerators; A Quality Assurance Tool. Medical Physics 28 (6): 1211, 2001.

Krintz, A., Followill, D., Melia, M., and Hanson W.F.: A Reanalysis of the Collaborative Ocular Melanoma Study Medium Tumor Trial Eye Plaque Dosimetry. Medical Physics 28 (6): 1192, 2001.

Tailor, R.C. and Hanson, W.F.: Estimated Absorbed -Dose Ratios "TG51/TG21" for Most Commonly Used Cylindrical and Parallel-Plate Ion Chambers over a Range of Photon and Electron Energies. Medical Physics 28 (6): 1192, 2001.

BenComo, J.A., Hernandez, N., and Hanson, W.F.: Problems and Shortcomings of the RPC Remote Monitoring Program of Institutions Dosimetry Data. Medical Physics 28 (6): 1212, 2001.

Hanson, W.F., Martin, B., Kuske, R., Arthur, D., Rabinovitch, R., White, J., Wilenzick, R. Harris, I., Tailor, R., and Davis, D.: Dose Specification and Quality Assurance of RTOG Protocol 95-17, A

192

Cooperative Group Study of Ir Breast Implants as Sole Therapy. Medical Physics 28 (6): 1297, 2001.

Lowenstein J, Roll J, Hanson W, Davis D, Lanciano R, Calkins A, Petereit, D, Varia M, and Ibbott, G, Radiotherapy Quality Assurance of Gynecologic Oncology Group (GOG) Protocol 165, A Cooperative Group Study of Carcinoma of the Cervix, Int. J. Rad. Onc., Bio., Phys., 54:2, 283, 2002.

Cho, S. H. and Ibbott, G.S. Reference Dosimetry Data for the Siemens Primus Linear Accelerator: Preliminary Results for Depth Dose and Output Factor, Medical Physics 29:1944, 2002.

Krintz, A., Hanson, W., Ibbott, G., and Followill, D., Verification of PLAQUE SIMULATOR dose distributions using radiochromic film. Medical Physics

29: 1220-1221, 2002.

Page 6 of15

Kry, S., Salehpour, M., Followill, D., Stovall, M., and Rosen, I., Secondary Dose from IMRT treatments. Medical Physics 29:1348, 2002.

Krintz, A., Hanson, W., Ibbott, G., and Followill, D., A reanalysis of the Collaborative Ocular Melanoma Study medium tumor trial eye plaque dosimetry. Int. J. Radiat. Oncol. Biol. Phys. 54: 329-330, 2002.

Duzenli, C., Baldock, C., Ibbott, G., Schreiner, L.J., Jordan, K., Oldham, M., Short, K., CCPM/COMP Symposium on Gel Dosimetry. Med. Phys. 29:1313, 2002.

Nath, R., Rivard, M., Coursey, B., DeWerd, Hanson, W., Huq, M., Ibbott, G., Williamson, J., Status of the American Association of Physicists in Medicine Radiation Therapy Committee’s Subcommittee on Low-Energy Interstitial Brachytherapy Source Dosimetry: Procedure for the Development of Consensus Single-Source Dose-Distribution. Med. Phys. 29:1349, 2002.

Beach, M.L., Ibbott, G.S., and Maryanski, M., Implementation of a Polymer Gel Dosimetry Insert for an Anthropomorphic Head Phantom used to Evaluate Head and Neck Intensity-Modulated Radiation Therapy. Med. Phys. 29:1942, 2002.

Cho, S.H. and Ibbott, G.S., Reference Photon Dosimetry Data for the Siemens Primus Linear Accelerator: Preliminary Results for Depth Dose and Output Factor. Med. Phys. 29:1944, 2002.

17.	Aguirre, J.F., Tailor, R., Ibbott, G., Stovall, M. Hanson,

W. TLD as a tool for remote verification of output for radiotherapy beams: 25 years of experience. Standards and Codes of Practice in Medical Radiation Dosimetry 2:191-9, 2002.

Ibbott, G., Beach, M., Maryanski, M. An anthropomorphic head phantom with a BANG® polymer gel insert for dosimetric evaluation of IMRT treatment delivery. Standards and Codes of Practice in Medical Radiation Dosimetry 2:361-8, 2002.

Ibbott, G., Nelson, A., Followill, D., Balter, P., Hanson,

W. An anthropomorphic head and neck phantom for evaluation of intensity modulated radiation therapy. Standards and Codes of Practice in Medical Radiation Dosimetry 2:209-17, 2002.

Izewska, J., Svensson, H., Ibbott, G. Worldwide QA networks for radiotherapy dosimetry. Standards and Codes of Practice in Medical Radiation Dosimetry 2:139-55, 2002.

Nelson, A., Balter, P., Hanson, W., Ibbott, G., Experience with an IMRT Head and Neck QA Phantom. Medical Physics, 29:1366, 2002.

Bencomo, Jose A., and Ibbott, Geoffrey S., Quality Assurance Measurements of Beam parameters of a Linear Accelerator using MOSFET Dosimeters: A Feasibility Study. Medical Physics, 29, (6): 1201, 2002.

Lowenstein J, Roll J, Ibbott G: To Plan or Not to Plan: Multiple High Dose Rate (HDR) Brachytherapy Insertations, Medical Physics, 30:1464, 2003.

Bayouth, J., Followill, D., Nelson, A., Ibbott, G., and Morrill, S., Dosimetric Effect of a Hip Prosthesis when Delivering Intensity Modulated Radiation Therapy (IMRT): A Phantom Study. Radiation and Oncology 68:S105, 2003.

Kry, S., Salehpour, M., Followill, D., Stovall, M., Rosen, I. Risk Assessment of Secondary Malignancies from IMRT Treatments, Medical Physics, 30 (6): 1130, 2003.

Heard, M., Ibbott, G., Followill, D. Characterizing Dose Distributions of Brachytherapy Sources Using Normoxic Gel (WIP), Medical Physics, 30 (6): 1333, 2003.

Esteban, J., Ibbott, G., Hanson, W., Horton, J., White, A., Cho, S. Energy Dependence of a TLD System for Characterizing Low Energy Brachytherapy Sources (WIP), Medical Physics, 30 (6): 1349, 2003.

Bencomo, J., Ibbott, G., Lee, S., Lii, M., Borges, J. Evaluation of Two Linear Ionization-Chamber-Arrays as Data Acquisition Systems for Quality Assurance of Photon and Electron Beams, Medical Physics, 30 (6): 1350, 2003.

Beach, M., Ibbott, G., Followill, D., Hanson, W., Bloch, C., Jackson, E., Tucker, S. Implementation of a Polymer Gel Dosimetry Insert for An Anthropomorphic Phantom Used to Evaluate Head and Neck Intensity-Modulated Radiation Therapy, Medical Physics, 30 (6): 1396, 2003.

30.	Rivard, M.J., Coursey, B.M., DeWerd, L.A., Hanson, W.F., Huq, M.S., Ibbott, G.S., Nath, R., Williamson,

J.F. Update of AAPM Task Group No. 43 Report -A Revised Protocol for Brachytherapy Dose Calculations, Medical Physics, 30 (6): 1431, 2003.

Rogers, D., DeWerd, L., Ibbott, G., Huq, M. Changes in Co-60 Air-Kerma Standards: The Rationale for Change and the Impact On Clinical Practice, Medical Physics, 30 (6): 1442, 2003.

Vassiliev, O., Liu, H.H., Dong, L, Antolak, J., Cho, S., Mohan, R. A Monte-Carlo-Based System for Dose Calculation and Dosimetry QA of Photon IMRT (WIP), Medical Physics, 30 (6): 1453, 2003.

Cho, S.H., Kim, C.H., Ibbott, G.S. Differences in the Results of MCNP Simulation for Low Energy Photon Sources May Be Due to the Choice of Photon Cross Section Libraries, Medical Physics, 30 (6): 1464, 2003.

Tailor, R., Esteban, J., Ibbott, G.S. TLD Measurements in Liquid Water of Dosimetry Characteristics of a New 125I Seed, Medical Physics, 30 (6): 1470, 2003.

Borges, J.A., Bencomo, J., Ibbott, G.S. A 3 Dimensional Gel Dosimetry Lung Equivalent (WIP), Medical Physics, 30 (6): 1478, 2003.

Tailor, R., Followill, D.S., Hernandez, N., Zhu, T.S., Ibbott, G.S. Quality-Assurance Check of Collimator and Phantom-Scatter Factors, Medical Physics, 30 (6): 1483, 2003.

Fisher, G., Followill, D., Tolani, N., Ibbott, G. The Accuracy of 3-D Inhomogeneity Photon Algorithms in Commercial Treatment Planning Systems Using a Heterogeneous Lung Phantom (WIP), Medical Physics, 30 (6): 1507, 2003.

Tailor, R., Hanson, W., Wells, N., Ibbott, G. Are Photon and Electron Beam Calibrations More Consistent with TG-51 Than with TG-21?, Medical Physics, 30 (6): 1350, 2003.

Rivard, M.J., Goetsch, S.J., Drzymala, R.E., Bourland, J.D., DeWerd, L.A., Gibbons, J.P., Ibbott, G.S., Kunugi, K.A., Moskvin, V., Walker, L.D. A Working Group for Improving Consistency of Quality Assurance, Treatment Planning, and Clinical Implementation for Gamma Knife® Stereotactic Radiosurgery. Accepted for oral presentation to GK2004 Meeting.

Fisher, G., Followill, D., Ibbott, G. The Accuracy of 3-D Inhomogeneity Photon Algorithms in Commercial Treatment Planning Systems Using a Heterogeneous Lung Phantom. Med. Phys. 31:1714, 2004.

Alvarez, P., Molineu, A., Hernandez, N., Followill, D., Balter, P., Hanson, W., Ibbott, G. Anthropomorphic Phantoms for Quality Assurance in Radiation Therapy. Med. Phys. 31:1732, 2004.

Halvorsen, P., Das, I., Freedman, D., Ibbott, G., Parsai, E., Rice, R., Robin, T., Thomadsen, B. Effective Peer Review for Clinical Radiation Oncology Physicists: Task Group 103's Preliminary Findings. Med. Phys. 31:1743, 2004.

Followill, D., Lowenstein, J., Ibbott, G. High-Energy Photon Standard Dosimetry Data: A Quality Assurance Tool. Med. Phys. 31:1782, 2004.

Gifford, K., Horton, J., Jackson, E., Steger, T., Heard, M., Mourtada, F., Lawyer, A., Ibbott, G. Verification of Monte Carlo Calculations Around a Fletcher Suit Delclos Ovoid with Radiochromic Film and Normoxic Polymer Gel Dosimetry. Med. Phys. 31:1808, 2004.

Molineu, A., Alvarez, P., Hernandez, N., Followill, D., Ibbott, G. Evaluation of 60 IMRT Irradiations of An Anthropomorphic H&N Phantom. Med. Phys. 31:1822-3, 2004.

Purdy, J., Palta, J., Ibbott, G. The Advanced Technology QA Consortium (ATC). Med. Phys. 31:1833, 2004.

BenComo, J., Stewart, B., Wells, N., Ibbott, G. Could MOSFET Detectors Substitute TLD Dosimeter as a Remote Monitoring Device of Megavoltage Beams Output? Med. Phys. 31:1875, 2004.

Esteban, J., Tailor, R., Cho, S., Ibbott, G. Energy Dependence of a New TLD System for the Characterization of Low Energy Brachytherapy Sources. Med. Phys. 31:1884-5, 2004. Med. Phys. 31:

Heard, M., Ibbott, G. Gel Dosimetry Technique for Measurements in High Dose Gradients. Med. Phys. 31:1888, 2004.

Cho, S., Vassiliev, O., Liu, H., Ibbott, G., Mohan, R. On the Choice of Measured Depth Dose Data for the Monte Carlo Modeling of the 18 MV Photon Beam: Shifted Vs. Unshifted. Med. Phys. 31:1906, 2004.

Urie, M., Ulin, K., Followill, D., Ibbott, G., Olch, A., Palta, J., Purdy, J. Results and Analysis by QARC of the IMRT Benchmark Required by the NCI for Participation in Clinical Trials. Med. Phys. 31:1915-6, 2004.

BenComo, J., Cho, S., Sun, T., Lee, S., Ibbott, G. Use of MOSFET Detectors to Verify Dose Calculations in An Anthropomorphic Breast Phantom. Med. Phys. 31:1926, 2004.

53.Molineu A, Alvarez P, Hernandez N, Followill DS, Ibbott GS. Evaluation of IMRT for Institutions Participating in NCI Sponsored Clinical Trials. International Journal of Radiation Oncology Biology Physics 60:S319, 2004.

Kry S, Titt U, Poenisch F, Followill D, Vassiliev O, Mohan R, Salehpour M. A Monte Carlo Simulation of Out-Of-Field Radiation From An 18-MV Beam. Medical Physics 32: 1889, 2005.

Molineu A, Hernandez N, Alvarez P, Followill D, Ibbott G. IMRT Head and Neck Phantom Irradiations: Correlation of Results with Institution Size. Medical Physics 32:1983-4, 2005.

Davidson S, Followill D, Ibbott G, Prado K. The Evaluation of Several Commercial IMRT Treatment Planning Systems Heterogeneity Dose Calculation Algorithms Using An Anthropomorphic Thorax Phantom. Medical Physics 32:1988, 2005.

Homann K, Gates B, Salehpour M, Followill D, Kirsner S, Buchholz T, White R, Prado K. Evaluation of the Dose Within the Abutment Region Between Tangential and Supraclavicular Fields for Various Breast Irradiation Techniques. Medical Physics 32: 1995, 2005.

Lowenstein J, Roll J, Davis C, Holguin P, Duong H, Followill D, Ibbott G. Credentialing Requirements for NSABP B-39 / RTOG 0413. Medical Physics 32:2020-1, 2005.

Alvarez P, Molineu A, Hernandez N, Followill D, Ibbott G. Evaluation of Doses Delivered by SBRT to the Lung of An Anthropomorphic Thorax Phantom. Medical Physics 32: 2043, 2005.

Followill D, Molineu A, McGary J, Ibbott G. Clinical Reference Dosimetry of a "Hi-Art II" Helical Tomotherapy Machine. Medical Physics 32:2089, 2005.

Followill D, Lowenstein J, Jhingran A, Roll J, Hernandez N, Ibbott G. The Radiological Physics Center’s Anthropomorphic Quality Assurance Phantom Family: New Developments. Medical Physics 32:2129, 2005.

Shoales J, Followill D, Ibbott G, Balter P, Tolani N. Development of An Independent Audit Device for Remote Verification of 4D Radiotherapy. Medical Physics 32:2128, 2005.

Followill D, Molineu A, McGary J, Hernandez N, Ibbott G. Evaluation of the TomoTherapy Planning Station Heterogeneity Correction Algorithm Using An Anthropomorphic Phantom. Medical Physics 32:2167, 2005.

Bencomo J, Macey D, Lawyer A. Verification of Dose Point Kernels for Ir-192 Brachytherapy. Medical Physics 32:1951, 2005.

Bencomo J, Weathers R, Stovall M, Ibbott G. Reference Electron Beam Dosimetry Data Set: A Preliminary Analysis. Medical Physics 32:1999, 2005.

Heard M, De La Mora A, Adamovics J, Ibbott G. Evaluation of a New 3D Polyurethane Dosimeter for IMRT Verification. Medical Physics 32:2167, 2005.

Vass H, Ibbott G. Comparison of PDR Iridium and LDR Cesium Through Monte Carlo Simulation. Medical Physics 32:1955, 2005.

Yoder R, Ibbott G. Needs in Ionizing Radiation Measurements and Standards. Medical Physics 32:2053, 2005.

Lowenstein J, Roll J, Ibbott G. Common Dosimetry Errors in Cervix Patients Treated with Brachytherapy on Clinical Trials. Medical Physics 32:2107, 2005.

Ibbott G. The Radiological Physics Center's QA Activities. Medical Physics 32:2153-4, 2005.

Briere TM, Tailor RC, Tolani NB, Prado KL, Lane RG, Woo SY, Ha CS, Gillin MT, Beddar AS. In Vivo Dosimetry Using Disposable MOSFET Dosimeters for Total Body Irradiation. Medical Physics 32:1996, 2005.

Schild SE, McGinnis WL, Graham D, Hillman S, Ibbott G, Northfelt D, Garces Y, Yee G, Bollinger J, Jett J. Results of a Phase I Trial of Concurrent Chemotherapy and Escalating Doses of Radiation for Unresectable Non-Small Cell Lung Cancer. International Journal of Radiation Oncology Biology Physics 63:S44, 2005.

Page 9 of15

72.Molineu A, Alvarez P, Hernandez N, Followill DS, Ibbott GS. Analysis of Errors Made During 138 IMRT Irradiations of an Anthropomorphic Phantom. International Journal of Radiation Oncology Biology Physics 63:S58, 2005.

Simon W, Kozelka J, Rose M, Liu C, Palta J, Dempsey J, Lynch B, Bayouth J, Pavord D, Ibbott G, Followill D. LINAC dosimetry: benchmark data set uncertainty. Med Physics 33:2118, 2006.

Kry S, Followill D, White R, Salehpour M. Uncertainty analysis of risk of secondary fatal malignancies from radiotherapy treatments including IMRT. Med Physics 33:2257, 2006.

Simon W, Kozelka J, Rose M, Liu C, Palta J, Dempsey J, Lynch B, Bayouth J, Pavord D, Ibbott G, Tailor R, Followill D. LINAC dosimetry: benchmark data set requirements. Med Physics 33:2118, 2006.

Lowenstein J, Davis C, Roll J, Harris I, Hall F, Followill D, Ibbott G. The credentialing process for the NSABP B-39/RTOG 0413 partial breast irradiation trial. Med Physics 33:2140, 2006.

77.	Alvarez P, Hernandez N, Followill D, Tailor R, Ibbott

G. Characterization of EBT versus MD55 gafchromic® films for relative dosimetry measurements. Med Physics 33:2217-18, 2006.

Hecox R, Gibbons J, Followill D, Ibbott G. Dose calculation accuracy in the presence of high-z materials using megavoltage CT for treatment planning. Med Physics 33:2087-8, 2006.

Howell R, Kry S, Followill D, Hertel N, Wang Z, Salehpour M. Investigation of MLC effects on secondary neutron spectra for Varian, Siemens, and Elekta. Med Physics 33: 2249-50, 2006.

Alvarez P, Molineu A, Hernandez N, Followill D, Ibbott G. Evaluation of heterogeneity corrections algorithms through the irradiation of a lung phantom. Med Physics 33:2214, 2006.

Followill D, Molineu A, Alvarez P, Ibbott G. The state of radiotherapy physics through the eyes of a quality auditor. Med Physics 33:2283, 2006.

Davidson S, Prado K, Ibbott G, Followill D. Hetergeneity dose calculation accuracy in IMRT using an anthropomorphic thorax phantom. Med Physics 33:2106, 2006.

83.	Followill D, Kry S, Salehpour M. Measurements of secondary radiation for electron and proton accelerators. Med Physics 33:2238-9, 2006.

Invited Articles

1.	Greven, K.M., Levenback, C., Chao, C.K.S., Delaney, T, Del Priore, G., Eifel, P., Erickson, B.A., Followill, D., Gaffney, D., Garcia, M., Gerszten, K., Grigsby, P., Henderson, R., Hricak, H., Hsu, J., Jhingrin, A., Kaye, A., Kudelka, A., Lukka, H., Mutch, D., Nag, S., Rotman, M., Shefter, T., Smith, W., Stehman, F., Souhami, L., Wenzel, L., Winter, K.A., and Wolfson,

A. “Gynecologic Cancer Working Group,” Int. J., Radiation Oncology Biol. Phys, Vol. 51 No. 3 Supplement 2, 58-59, 2001.

Michalski, J., Purdy, J.A., Gaspar, L., Souhami, L., Ballow, M., Bradley, J., Chao, C.K.S., Crane, C., Eisbruch, A., Followill, D., Forster, K., Fowler, J., Gillin, M.T., Graham, M.L. Harms, W.B., Huq, S., Kline, R.W., Langer, M., Mackie, T.R., Mukherji, S., Podgorsak, E.B., Roach, M., Ryu, J., Sandler, H., Schultz, C.J., Schell, M., Verhey, L.J. Vicini, F., and Winter, K.A., “Image-Guided Radiation Therapy Committee”, Int. J., Radiation Oncology Biol. Phys, Vol. 51 No. 3 Supplement 2, 60-65, 2001.

Cho, S. H. Dosimetry of brachytherapy sources: Review of the AAPM TG-43 formalism, Proceedings of the 3rd Korea-Japan Joint Meeting on Medical Physics and the 2nd Asia Oceania Congress of Medical Physics 2002, ISSN 1589-6691, September 2002.

Ibbott GS. The medical physics consult -gel dosimetry. J of the Amer College of Radiol 3:144-6, 2006.

Letters to the Editor/Newsletters

Ibbott, G.S. and Nath, R.: “Dose-rate constant for Imagyn 125I brachytherapy source. Medical Physics 28(4), April 2001.

Ibbott, G.S., Meigooni, A.S., and Gearheart, D.M.: Monte Carlo determination of dose rate constant,” Medical Physics 29, 1637-1638, 2002.

Rivard, M.J., Coursey, B.M., DeWerd, L.A., Hanson, W.F., Huq, M.S., Ibbott, G.S., Nath, R., Williamson, J.F.: Comment on “Let’s abandon

geometry factors other than that of a point source in brachytherapy dosimetry,” Medical Physics 29, 1917-1918, 2002.

Followill, D.S., Lowenstein, J.R., and Ibbott, G, Quality Assurance: It’s Here to Stay, Newsletter American Association of Physicists in Medicine, Vol. 28, No. 1, January/February 2003.

Tailor R, Hanson W, Ibbott G. Response to Comment on Calculated absorbed-dose ratios, TG51/TG21, for most widely used cylindrical and parallel-plate ion chambers over a range of photon and electron energies. Letter to Editor, Medical Physics, 30:478-480,2003.

Olch, A., Followill, D., Thomadsen, B., Wichman, B., Rosemark, P., and Hunt, M., In regard to Urie et al.: Current calibration, treatment, and treatment planning techniques among institutions participating in the children'’ oncology group. Int. J., Radiation Oncology Biol. Phys, Vol. 56 No. 4:1209-1210, 2003.

7.	Palta, J.R., Deye, J.A., Ibbott, G.S., Purdy, J.A., Urie,

M.M. Credentialing of Institutions for Use of IMRT in Clinical Trials. In regard to Engler and Rivard: Evaluation of the NCI IMRT Benchmark for clinical trials. IJROBP 57:S260, 2003.

Palta JR, Deye JA, Ibbott GS, Purdy JA, Urie MM. Credentialing of Institutions for IMRT in Clinical Trials. Int. J. of Radiation Oncology, Biol. Phys. 59:1257-62, 2004.

Nag S, Cardenas H, Chang S, Das IJ, Ibbott GS, Thomadsen B, Varia M. Reply to GEC-ESTRO-GYN letter. Int. J. of Radiation Oncology, Biol. Phys. 62:295-6, 2005.

Nag S, Cardenas H, Chang S, Das IJ, Ibbott GS, Lowenstein J, Roll J, Thomadsen B, Varia M. Reply to Narayan et al regarding ROB-D-05-00575. Int. J. of Radiation Oncology, Biol. Phys. (in press), 2005.

Butler WM, Huq MS, Li Z, Thomadsen BR, DeWerd LA, Ibbott GS, Mitch MG, Nath R, Rivard MJ, Williamson JF, Yue NJ, Zaider M. Third party brachytherapy seed calibrations and physicist responsibilities. Medical Physics 33:247-8, 2006.

Manuals, Teaching Aids, Other Publications

Digital Imaging Technology Applications in Mammography in CD of proceedings of the 2nd meeting of ALFIM in Caracas Venezuela, October 2001.

Bencomo, J.A., Hernandez, N., and Hanson, W.F. Problems and Shortcomings of the RPC Remote Monitoring Program of Institutions Dosimetry Data. AAPM 43rd Annual Meeting, Salt Lake City, Utah, July 22-26, 2001.

Lowenstein, J., Educational video on how to utilize the TG-51 Calibration Protocol, 2001.

Ibbott, G., Manual on Radiation Protection in Hospitals and General Practice, ed. by Gerald P. Hanson, to be published by WHO, 2001.

Bencomo, Jose A., and Ibbott, Geoffrey S., Quality Assurance Measurements of Beam parameters of a Linear Accelerator using MOSFET Dosimeters: A feasibility Study. AAPM 44th Annual Meeting, Montreal, Ontario, Canada, June 13-18, 2002.

Aguirre, F., Physical aspects of quality assurance in radiotherapy recommendation from a panel of experts from the IAEA within the activation of the Regional Program ARCAL XXX (to be published).

Aguirre, F. and Ibbott, G., Lecture materials for "Prevention of Accidental Exposure", IAEA Division of Radiation and Waste Safety, 2003.

Followill, D., Lowenstein, J., and Ibbott, G., Quality Assurance: It’s here to stay. AAPM Newsletter, Jan/Feb, 2003.

Olch, A., Kline, R., Ibbott, G., et. al, Quality Assurance for Clinical Trials: A Primer for Physicists. Prepared by AAPM Subcommittee on QA for Clinical Trials. AAPM Report No. 86, 2004.

Book Chapters

BenComo, José A., Introduction to Radiation Protection for Medical Physicists, Text Book CD-ROM (under development).

Dixon RL, Butler PF, Sobol WT. What every medical physicist should know about the JCAHO standards. In: Accreditation Programs and the Medical Physicist, Medical Physics Publishing, Madison, pp. 139-73, 2001.

Balter S, Chan RC, Shope TB. Professional roles in VBT. In: Intravascular Brachytherapy Fluoroscopically Guided Interventions, Medical Physics Publishing, Madison, pp. 531-44, 2002.

Galvin JM, Ibbott GS. Commissioning and accreditation of a stereotactic body radiation therapy program. In: Stereotactic Body Radiation Therapy, Lippincott Williams & Wilkins, Philadelphia, pp. 85-93, 2005.

Webster JG. Radiation dosimetry: 3-dimensional. In: Encyclopedia of Medical Devices and Instrumentation, John Wiley & Sons, Hoboken, pp. 481-500, 2006.

Books (edited and written)

1.	Radiation Therapy Physics, 2nd edition, W.R. Hendee,

G. S. Ibbott, Mosby Year-Book Publishers, Philadelphia, 1996.

Principles of Digital Mammography 1st Ed. (University Autonomous of Mexico City, Mex. (under Development), 2002.

Radiation Therapy Physics, 3rd edition, W.R. Hendee,

G.S. Ibbott, and E.G. Hendee, John Wiley and Sons, New York, 2004.

PRESENTATIONS INTERNATIONAL ACTIVITIES

Geoffrey Ibbott attended the DOSGEL Meeting, Sherbrooke, Canada, August 7-10, 2006.

Geoffrey Ibbott attended the ACR CSC Meeting, Jackson Hole, WY, August 18-21, 2006.

Geoffrey Ibbott attended the IEC Meeting, Delft, Holland, October 2-6, 2006.

Geoffrey Ibbott attended the ESTRO Meeting, Leipzig, Germany, October 8-12, 2006.

Geoffrey Ibbott presented in Seoul, Korea, October 18-19, 2006.

Geoffrey Ibbott presented at Duke University, Durham, NC, October 27, 2006.

VISITS TO INSTITUTIONS

1.	Geoffrey Ibbott performed radiological physics measurements and reviewed patient dosimetry at the Hahnemann University, Philadelphia, PA, August 14-15, 2006.

Geoffrey Ibbott performed radiological physics measurements and reviewed patient dosimetry at the University of Florida, Jacksonville, FL, August 17-18, 2006.

David Followill performed radiological physics measurements and reviewed patient dosimetry at the Juravinski Cancer Center, Hamilton, Ontario, Canada, September 25-28, 2006.

Jessica Lowenstein Leif performed radiological physics measurements and reviewed patient dosimetry at the Juravinski Cancer Center, Hamilton, Ontario, Canada, September 25-28, 2006.

Susannah Lazar performed radiological physics measurements and reviewed patient dosimetry at the Mary Bird Perkins Cancer Center, Baton Rouge, LA, September 28-29, 2006.

Francisco Aguirre performed radiological physics measurements and reviewed patient dosimetry at the Veterans Administration Hospital, Miami, FL, September 29-October 2, 2006.

Susannah Lazar performed radiological physics measurements and reviewed patient dosimetry at the Mary Bird Perkins Cancer Center, Baton Rouge, LA, October 19-21, 2006.

MEETINGS ATTENDED

(August 1, 2006 -October 31, 2006)

Geoffrey Ibbott attended the AAPM Meeting, Orlando, FL, July 28-August 3, 2006.

David Followill attended the AAPM Meeting, Orlando, FL, July 29-August 3, 2006.

Malcolm Heard attended the AAPM Meeting, Orlando, FL, July 29-August 3, 2006.

Respectfully submitted, Page 12 of 15

Jessica Lowenstein Leif attended the AAPM Meeting, Orlando, FL, July 29-August 3, 2006.

Jose Bencomo attended the AAPM Meeting, Orlando, FL, July 29-August 3, 2006.

Joye Roll attended the AAPM Meeting, Orlando, FL, July 29-August 3, 2006.

Elizabeth Siller attended the AAPM Meeting, Orlando, FL, July 29-August 3, 2006.

Nadia Hernandez attended the AAPM Meeting, Orlando, FL, July 29-August 3, 2006.

Ryan Hecox attended the AAPM Meeting, Orlando, FL, July 29-August 3, 2006.

Nathan Wells attended the AAPM Meeting, Orlando, FL, July 29-August 3, 2006.

Scott Davidson attended the AAPM Meeting, Orlando, FL, July 31-August 3, 2006.

Geoffrey Ibbott attended the NCCTG Meeting, Rochester, MN, September 26-28, 2006.

David Followill attended the COG Meeting, Los Angeles, CA, October 4-6, 2006.

Francisco Aguirre attended the SWOG Meeting, Seattle, WA, October 5-8, 2006.

Irene Harris attended the SWOG Meeting, Seattle, WA, October 5-8, 2006.

Geoffrey Ibbott attended the CIRMS Meeting, Gaithersburg, MD, October 22-25, 2006.

Geoffrey S. Ibbott, Ph.D.

EXPENDITURES OF THE RADIOLOGICAL PHYSICS CENTER (RPC Grant and the Advanced Technology Subcontract)

August 1, 2006 -October 31, 2006

PERSONNEL (salaries, fringe benefits): $730,809

8 Physicists, 1 Supervisor of Quality Assurance Dosimetry Services, 4 Dosimetrists, 1 Informatics Manager, 1 Database Administrator, 1 Programmer Analyst I, 1 Radiological Physics Supervisor, 5 Radiological Physics Technicians, 1, Coordinator of Research Data, 1 Department Administrator, 1 Sr. Administrative Assistant, 2 Administrative Assistants, 1 Procurement Technician, 1 Sr. Secretary, and 3 Graduate Research Assistants.

______________________________________________________________________________

TRAVEL Visits:

| Geoffrey Ibbott Geoffrey Ibbott David Followill Jessica Lowenstein Susannah Lazar Francisco Aguirre Susannah Lazar | Hahnemann University, Philadelphia, PA University of Florida, Jacksonville, FL Juravinski Cancer Center, Hamilton, Ontario, Canada Juravinski Cancer Center, Hamilton, Ontario, Canada Mary Bird Perkins Cancer Center, Baton Rouge, LA Veterans Administration Hospital, Miami, FL Mary Bird Perkins Cancer Center, Baton Rouge, LA |  | $ 674.00 $ 843.10 $1,594.57 $ 895.63 $ 278.07 $1,043.95 $ 278.07 |
| --- | --- | --- | --- |
|  |  | Total | $ 5607.39 |

Meetings:

| Geoffrey Ibbott David Followill Malcolm Heard Jessica Lowenstein Jose Bencomo Joye Roll Elizabeth Siller Nadia Hernandez Ryan Hecox Nathan Wells Scott Davidson Geoffrey Ibbott David Followill Francisco Aguirre Irene Harris Geoffrey Ibbott | AAPM, Orlando, FL AAPM, Orlando, FL AAPM, Orlando, FL AAPM, Orlando, FL AAPM, Orlando, FL AAPM, Orlando, FL AAPM, Orlando, FL AAPM, Orlando, FL AAPM, Orlando, FL AAPM, Orlando, FL AAPM, Orlando, FL NCCTG, Rochester, MN COG, Los Angeles, CA SWOG, Seattle, WA SWOG, Seattle, WA CIRMS, Gaithersburg, MD | $1,558.34 $ 783.76 $ 430.19 $ 758.78 $ 744.68 $ 641.26 $ 266.55 $ 695.80 $ 350.00 $ 966.35 $ 490.68 $ 517.70 $ 463.23 $ 970.38 $ 521.40 $1,430.10 |
| --- | --- | --- |
|  |  | Total $ 11,589.20 |

| CONSULTANTS | $-0 |
| --- | --- |
| SUPPLIES: Office supplies, laboratory and record keeping, TLD, TLD supplies, software, equipment, etc. | $11,302 |
| OTHER EXPENSE: Postage, telephone, reprints, copying, computer fees, equipment repair, registration fees, tuition, freight/delivery, etc. | $6,237 |

| RPC Report to the AAPM Therapy Physics Committee | Page 14 of 15 |
| --- | --- |
| Report No. 125 |  |
| SPACE RENTAL: | $59,574 |
| Total Expenditures August 1, 2006 -October 31, 2006 | $825,118.59 |
| Indirect costs @ 26% | $214,530.83 |
| TOTAL | $1,039,649.42 |

| RPC Report to the AAPM Therapy Physics Committee |  | Page 15 of 15 |
| --- | --- | --- |
| Report No. 125 |  |  |
| RPC Report to TPC October 2006 |  |  |
| Clinical Study Groups | Office Reviewing | Special Projects |
|  | Patient Records |  |
|  |  | Radiotherapy manual |
| Gynecologic Oncology Group GOG | RPC | Electronic Transfer of Patient Records Image Based Treatment Planning |
|  |  | IMRT Guidelines |
|  |  | Defining Treatment violations |
| National Surgical Adjuvant Breast and Bowel Project NSABP | RPC | IMRT Guidelines Breast phantom Partial Breast RT Credentialing |
| North Central Cancer Treatment Group NCCTG | RPC | Rapid Review of Lung Study 3D CRT credentialing Stereotactic Phantom |
|  |  | IMRT H&N Phantom |
|  |  | Prostate Phantom |
|  |  | Prostate Implant |
|  |  | Credentialing LDR/HDR |
| Radiation Therapy Oncology Group RTOG | RTOG/RPC | Stereotactic Head Phantom Lung Phantom 4D Liver Phantom |
|  |  | Cervix HDR/LDR protocol compliance |
|  |  | Patient Calculation Program |
|  |  | Partial Breast RT Credentialing |
|  |  | IMRT Benchmark Case |
| Southwest Oncology Group SWOG | QARC | 3D Benchmark Case |
|  |  | RPC Institution List |
| Clinical Trial Support Unit CTSU | QARC, RPC, RTOG | RTF Numbers TLD Monitoring |
|  |  | Review RT Facility Questionnaire |
|  |  | Participate | in | the | development | of |
| American College of Radiology Imaging Network ACRIN | N/A | guidelines for quality assurance institution participating in ACRIN | of |
|  |  | CT Dose Measurements |
| American College of Surgeons Oncology Group ACOSOG | QARC | RPC Institution List |
| Cancer and Acute Leukemia Group B CALGB | QARC | TRUS Prostate Approval Collaboration |
| Children’s Oncology Group COG (POG, CCG, WILMS, RHABDO) | QARC | 3D Benchmark Case IMRT Benchmark Case/Phantom CT/MRI Fusion Benchmark |
| Eastern Cooperative Oncology Group ECOG | QARC |  |